## Sheet1
ลำดับ (Item No.) | รายละเอียดเครื่องมือ (Description) | ช่วงการวัด (Range) | ความสามารถการวัด (CMC / Uncertainty) | ราคาค่าบริการ (Price) | หมายเหตุ (Remark) | รหัสบริการ (Code) | ฝ่ายงานจริง (Actual Department) | ราคาตัวเลข (Price Numeric) | ฝ่ายงานที่ทำนาย (Predicted Department) | รหัสกลุ่มคลัสเตอร์ (Cluster ID) | คำสำคัญของกลุ่มคลัสเตอร์ (Cluster Keywords)
1 | Iodine Stabilized He-Ne laser | 633 nm | 5.00E-11 | 36,000/piece | Beat measurement mehod with Iodine stabilized He-Ne laser | 08011 - 10101 | Dimensional Metrology Department | 36000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2 | Stabilized He-Ne laser | 633 nm | 1.00E-09 | 22,530/piece | Beat measurement mehod with Iodine stabilized He-Ne laser | 08011 - 10201 | Dimensional Metrology Department | 22530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
3 | Stabilized He-Ne laser | 500 - 1000 nm and 1560 nm | 1.00E-12 | 45,000/piece | Beat measurement mehod with optical comb | 08011 - 10202 | Dimensional Metrology Department | 45000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
4 | Stabilized He-Ne laser | 532 nm | 4.00E-08 | 22,530/piece | Beat measurement mehod with Iodine stabilized Nd:YAG laser | 08011 - 10203 | Dimensional Metrology Department | 22530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
5 | Dual frequency stabilized He-Ne laser | 633 nm | 1.00E-09 | 25,000 | Beat measurement mehod with Iodine stabilized He-Ne laser | 08011 - 10204 | Dimensional Metrology Department | 25000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
6 | Nonstabilized Laser | 633 nm | 1.00E-07 | 22,530/piece | Beat measurement mehod with Iodine stabilized He-Ne laser | 08011 - 10301 | Dimensional Metrology Department | 22530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
7 | Nonstabilized Laser | 350 - 1100 nm | 1.70E-07 | 21,280/piece | Direct measurment method by wavelength meter | 08011 - 10401 | Dimensional Metrology Department | 21280 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
8 | Wavelength meter | 532 - 1064 nm | 4.00E-08 | 25,000/piece | Direct measurment method by stabilized light source | 08011 - 10503 | Dimensional Metrology Department | 25000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
12 | Rectangular Gauge Block, Material : Chromium Carbide | 0.5 mm to 125 mm | Q[26 nm, 0.27E-6 L] | 2,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10101 | Dimensional Metrology Department | 2000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
11 | Rectangular Gauge Block, Material : Tungsten Carbide | 0.5 mm to 125 mm | Q[26 nm, 0.18E-6 L] | 2,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10101 | Dimensional Metrology Department | 2000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
9 | Rectangular Gauge Block, Material : Steel | 0.5 mm to 125 mm | Q[26 nm, 0.35E-6 L] | 2,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10101 | Dimensional Metrology Department | 2000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
10 | Rectangular Gauge Block, Material : Ceramic | 0.5 mm to 125 mm | Q[26 nm, 0.29E-6 L] | 2,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10101 | Dimensional Metrology Department | 2000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
16 | Rectangular Gauge Block, Material : Chromium Carbide | < 0.5 mm | Q[26 nm, 0.27E-6 L] | 3,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10102 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
13 | Rectangular Gauge Block, Material : Steel | < 0.5 mm | Q[26 nm, 0.35E-6 L] | 3,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10102 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
14 | Rectangular Gauge Block, Material : Ceramic | < 0.5 mm | Q[26 nm, 0.29E-6 L] | 3,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10102 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
15 | Rectangular Gauge Block, Material : Tungsten Carbide | < 0.5 mm | Q[26 nm, 0.18E-6 L] | 3,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10102 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
17 | Rectangular Gauge Block, Material : Steel, Ceramic, Tungsten Carbide, Chromium Carbide - Variation in length | 0.5 mm to 125 mm | 0.020 um | 2,200/piece | Optical Interferometry method, Variation (fo, fu) in length measurement. Measuring faces inspection without lapping. | 08021 - 10103 | Dimensional Metrology Department | 2200 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
18 | Rectangular Gauge Block, Material : Steel, Ceramic, Tungsten Carbide, Chromium Carbide - Flatness | 0.5 mm to 125 mm | 35 nm | 2,200/piece | Optical Interferometry method, Flatness measurement | 08021 - 10104 | Dimensional Metrology Department | 2200 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
19 | Rectangular Gauge Block, according to Euramet cg-2/v.02 (03/2011) | 0.5 mm to 100 mm | 26 nm | 2,500/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 10201 | Dimensional Metrology Department | 2500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
20 | Rectangular and Square Gauge Block (Metric), Material : Steel, Tungsten Carbide, Chromium Carbide - Deviation from central length | 0.5 mm to 100 mm | Q[50 nm, 0.50E-06 L] | 750/piece | Comparison method, Dlc, Deviation from central length , Variation (fo, fu) Measuring faces inspection without lapping | 08021 - 10301 | Dimensional Metrology Department | 750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
22 | Rectangular and Square Gauge Block (Metric), Material : Ceramic - Deviation from central length | 0.5 mm to 100 mm | Q[70 nm, 0.50E-06 L] | 750/piece | Comparison method, Dlc, Deviation from central length , Variation (fo, fu) Measuring faces inspection without lapping | 08021 - 10301 | Dimensional Metrology Department | 750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
21 | Rectangular and Square Gauge Block (Metric), Material : Steel, Tungsten Carbide, Chromium Carbide - Deviation from central length | >100 mm to 125 mm | Q[50 nm, 0.50E-06 L] | 3,030/piece | Comparison method, Dlc, Deviation from central length , Variation (fo, fu) Measuring faces inspection without lapping | 08021 - 10302 | Dimensional Metrology Department | 3030 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
23 | Rectangular and Square Gauge Block (Metric), Material : Ceramic - Deviation from central length | >100 mm to 125 mm | Q[70 nm, 0.50E-06 L] | 3,030/piece | Comparison method, Dlc, Deviation from central length , Variation (fo, fu) Measuring faces inspection without lapping | 08021 - 10302 | Dimensional Metrology Department | 3030 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
24 | Rectangular and Square Gauge Block (Metric), Material : Steel, Ceramic, Tungsten Carbide, Chromium Carbide - Difference in central length | 0.01 - 25 mm | 47 nm | 1,905/pair | Direct Measurement by gauge block comparator, Difference in central length, variation in length and measuring face inspection without lapping. | 08021 - 10303 | Dimensional Metrology Department | 1905 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
26 | Rectangular and Square Gauge Block (inch), - Deviation from central length Material : Ceramic | 0.005 inch to 4 inch | Q[70 nm, 0.50E-06 L] | 750/piece | Comparison method, Dlc, Deviation from central length , Variation (fo, fu) Measuring faces inspection without lapping | 08021 - 10501 | Dimensional Metrology Department | 750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
25 | Rectangular and Square Gauge Block (inch), - Deviation from central length Material : Steel, Tungsten Carbide, Chromium Carbide | 0.005 inch to 4 inch | Q[50 nm, 0.50E-06 L] | 750/piece | Comparison method, Dlc, Deviation from central length , Variation (fo, fu) Measuring faces inspection without lapping | 08021 - 10501 | Dimensional Metrology Department | 750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
27 | Long Gauge Block (Metric Rectangular, Square) Material : Steel | 125 mm to 500 mm | Q[50 nm, 0.80E-6 L] | 3,380/piece | Comparison Method, Dlc, Deviation from central length, Variation (fo, fu) Measuring faces inspection without lapping. | 08021 - 10601 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
28 | Long Gauge Block (Metric Rectangular, Square) Material : Steel | 600 mm to 1000 mm | Q[50 nm, 0.80E-6 L] | 6,750/piece | Comparison Method, Dlc, Deviation from central length, Variation (fo, fu) Measuring faces inspection without lapping. | 08021 - 10603 | Dimensional Metrology Department | 6750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
29 | Length Bar | Up to 800 mm | Q[50 nm, 0.80E-6 L] | 3,380/piece | Comparison Method, Dlc, Deviation from central length, Variation (fo, fu) Measuring faces inspection. | 08021 - 10701 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
30 | Ball Bar / Nest Bar | up to 800 mm | 1.2 ?m | 3,780/piece | Direct measurement by ULM and laser interferometer | 08021 - 10703 | Dimensional Metrology Department | 3780 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
31 | Long Gauge Block (Metric Rectangular) Material : Steel | >150 mm to 500 mm | Q[20 nm, 0.44E-6 L] | 10,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection and lapping. Both side wringing | 08021 - 10801 | Dimensional Metrology Department | 10000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
32 | Long Gauge Block (Metric Rectangular) Material : Steel | 600 mm to 1000 mm | Q[20 nm, 0.44E-6 L] | 15,000/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection and lapping. Both side wringing | 08021 - 10802 | Dimensional Metrology Department | 15000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
33 | Square Gauge Block Material : Steel | 0.5 mm to 125 mm | Q[28 nm, 0.35E-6 L] | 1,730/piece | Optical Interferometry method, Deviation from central length Measuring faces inspection without lapping. Both side wringing | 08021 - 11001 | Dimensional Metrology Department | 1730 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
34 | Step Master ; Material : Steel | Up to 300 mm | 0.028 mm | 12500 | Optical Interferometry method, Length difference | 08021 - 11101 | Dimensional Metrology Department | 12500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
35 | Step Master ; Material : Ceramic | Up to 300 mm | 0.030 mm | 12500 | Optical Interferometry method, Length difference | 08021 - 11102 | Dimensional Metrology Department | 12500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
194 | Gauge Block Comparator | 0.5 mm to 100 mm | 37 nm | 25,000/piece | Cal. by 11 Gauge Blocks 0.5 to 100 mm Euramet cg-2/v.02 (03/2011) | 08021 - 30801 | Dimensional Metrology Department | 25000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
184 | Universal Length Measuring Machine (ULM) | Up to 500 mm | Q[0.050 um, 0.80E-6 L] | 30,000/piece (10 points) + 1,000/point | Direct measurement with gauge block | 08021 - 30901 | Dimensional Metrology Department | 30000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
185 | Universal Length Measuring Machine (ULM) | >500 mm to 1000 mm | Q[0.060 um, 0.80E-6 L] | 30,000/piece (10 points) + 1,000/point | Direct measurement with gauge block | 08021 - 30901 | Dimensional Metrology Department | 30000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
186 | Universal Length Measuring Machine (ULM) | Up to 2000 mm | Q[0.040 um, 0.40E-6 L] | 40,000/piece (10 points) + 1,000/point | Direct measurement with Laser Interferometer | 08021 - 30902 | Dimensional Metrology Department | 40000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
36 | Autocollimator (Analog, Digital) | ± 1000" | 0.2" | 25,000/piece | Calibration by standard autocollimator | 08031 - 10101 | Dimensional Metrology Department | 25000 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
37 | Autocollimator (Analog, Digital) | ± 1500" | 0.15" | 25,000/piece | Calibration by Self calibration angle system | 08031 - 10103 | Dimensional Metrology Department | 25000 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
38 | Rotary Encoder | 0 - 360 degree | 0.06" | 25,000/piece | Calibration by self calibration angle system (SCMS-127) | 08031 - 10201 | Dimensional Metrology Department | 25000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
39 | Rotary Encoder | 0 - 360 degree | 0.11" | 25,000/piece | Calibration by self calibration angle system (SCMS-107) | 08031 - 10202 | Dimensional Metrology Department | 25000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
40 | Rotary Encoder | 0 - 360 degree | 7" | 25,000/piece | Calibration by self calibration angle system | 08031 - 10203 | Dimensional Metrology Department | 25000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
41 | Polygon | 3 to 72 faces | 0.2" | 1,000/face | Calibration by one autocollimator | 08031 - 10301 | Dimensional Metrology Department | 1000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
42 | Polygon | 3 to 72 faces | 0.25" | 750/face | Calibration by two autocollimator | 08031 - 10305 | Dimensional Metrology Department | 750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
43 | Polygon | 3 to 72 faces | 0.2" | 1,000/face | Calibration by self calibration rotray encoder system | 08031 - 10309 | Dimensional Metrology Department | 1000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
44 | Angle Gauge Block | up to 90o | 0.33" | 1,500/piece | Calibration by autocalimator | 08031 - 10401 | Dimensional Metrology Department | 1500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
45 | Indexing Table | 0o to 360o | 0.2" | 1,000/position | Calibration by one autocollimator | 08031 - 10501 | Dimensional Metrology Department | 1000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
46 | Electronic Inclinometer | ±1000" | 0.35" | 15,000/range | Calibration by autocollimator | 08031 - 10901 | Dimensional Metrology Department | 15000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
47 | Electronic Level | ±60o | 0.002o | 5,000 | Calibration by rotary encorder | 08031 - 10902 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
195 | Precision Level | Full range | 5.4 mm/m | 5,630/piece | Calibration by autocollimator | 08031 - 11001 | Dimensional Metrology Department | 5630 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
48 | Cylindrical Square | Up to 400 mm | 2.3 ?m | 5,000/side | Calibration by Master Square | 08031 - 11101 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
49 | Standard Square | Up to 400 mm | 2.3 µm | 5,000/side | Calibration by Master Square and Measuring Probe | 08031 - 11201 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
50 | Squareness Tester | Up to 400 mm | 2.0 µm | 5,000 | Calibration by Master Square | 08031 - 11301 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
51 | Precision Square | Up to 400 mm | 2.5 ?m | 5,000 | Calibration by Master Square & Gauge Blocks | 08031 - 11401 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
196 | Bevel protractor | 0° to 360° | 1.3' | 5000 | Cal. By angle gauge block | 08031 - 11501 | Dimensional Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
52 | Square and Squareness Tester | Up to 1000 mm | 0.4 µm | 10,000 (For up to 10 points) + 1,000/point | Calibration by using reversal method | 08031 - 11501 | Dimensional Metrology Department | 10000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
53 | Small Angle Generator | ?5,000 ?m/m | 2 ?m/m | 15000 | Calibration by autocollimator | 08031 - 11901 | Dimensional Metrology Department | 15000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
54 | Sine Bar | up to 45o | 0.0021o | 5000 | Calibration by Gauge Blocks and Angle Gauge Blocks | 08031 - 12001 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
55 | Straightedge | up to 1000 mm | 1.0 ?m | 5,000 | Calibration straighness measuring machine | 08031 - 12201 | Dimensional Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
56 | Straightness Measuring Machine | up to 2,500 mm | (1.6E-3 L) ?m | 5,000 | Calibration by standard autocollimator | 08031 - 12202 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
57 | Straightness Measuring Machine and Straightedge | up to 3000 mm | 0.7 ?m | 10,000/m | Calibration by using reversal method | 08031 - 12203 | Dimensional Metrology Department | 10000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
197 | Surface plate | 250x160 mm to 2,500x1,600 mm | 0.5 to 5.5 um | 15,000 | Cal by Inclinometer | 08031 - 12301 | Dimensional Metrology Department | 15000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
187 | Rotary encoder | 0-360 degree | 5" | 30,000 | Comparative measurement using reference encoder | 08031 - 30201 | Dimensional Metrology Department | 30000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
58 | Plain Ring Gauge | 0.1 to 100 mm | Q[0.37 µm, 1.3E-06 D] | 2,410/piece | Calibration by Small Internal Diameter Measurement (IDM) | 08041 - 10101 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
59 | Plain Ring Gauge | 0.5 mm to 300 mm | Q[0.25 µm, 1.8E-06 D] | 4,410/piece | Compare measurement method with setting plain ring gauge | 08041 - 10102 | Dimensional Metrology Department | 4410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
60 | Plain Ring Gauge (Hole) | 2.5 mm - 160 mm | 1.5 ?m | 5,000/piece | Contact scanning with measuring probe | 08041 - 10103 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
61 | Go - No Go Plain Ring Gauge | 0.1 to 100 mm | Q[0.37 µm, 1.3E-06 D] | 2,410/side | Calibration by Small Internal Diameter Measurement (IDM) | 08041 - 10201 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
62 | Taper Plain Ring Gauge | Pitch 0.45 mm to 8 mm | 2.0 µm | 2,410/piece | Cal. by Universal length measuring Machine | 08041 - 10301 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
63 | Go - No Go Taper Plain Ring Gauge | Pitch 0.45 mm to 8 mm | 2.0 µm | 2,410/side | Cal. by Universal length measuring Machine | 08041 - 10401 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
64 | Plain Plug Gauge | 0.1 to 300 mm | Q[0.20 µm, 1.0E-06 D] | 2,410/piece | Cal. by Universal length measuring Machine | 08041 - 10501 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
65 | Plain Plug Gauge (Shaft) | 1 mm - 150 mm | 1.5 ?m | 5,000/piece | Contact scanning with measuring probe | 08041 - 10502 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
66 | Go-No Go Plain Plug Gauge | 0.1 to 300 mm | Q[0.20 µm, 1.0E-06 D] | 2,410/side | Cal. by Universal length measuring Machine | 08041 - 10601 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
198 | Pin Gauge | 0.1 to 1 mm | Q[0.20 µm, 1.0E-06 D] | 2,410/piece | Cal. by Universal length measuring Machine | 08041 - 10701 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
199 | Pin Gauge | 1 to 100 mm | Q[0.20 µm, 1.0E-06 D] | 2,410/piece | Cal. by Universal length measuring Machine | 08041 - 10702 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
200 | Pin Gauge | 100 to 300 mm | Q[0.20 µm, 1.0E-06 D] | 2,410/piece | Cal. by Universal length measuring Machine | 08041 - 10703 | Dimensional Metrology Department | 2410 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
67 | Sphere | dia. Up to 300 mm | Q[0.20 µm, 1.0E-06 D] | 2,410/piece | Cal. by Universal length measuring Machine | 08041 - 10801 | Dimensional Metrology Department | 2410 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
68 | 3-Wire Unit | All sizes | Q[0.20 µm, 1.0E-06 D] | 2,820/set | Cal. by Universal length measuring Machine | 08041 - 10901 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
69 | Taper Plain Plug Gauge | from pitch 0.45 - 8 mm | 2.0 µm | 3,500/piece | Cal. by Universal length measuring Machine and T-Ball Sttylus | 08041 - 11501 | Dimensional Metrology Department | 3500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
70 | Parallel Thread Ring Gauge | from pitch 0.45 - 8 mm | 1.5 µm | 3,500/piece | Cal. by Universal length measuring Machine and T-Ball Sttylus | 08041 - 11701 | Dimensional Metrology Department | 3500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
71 | Parallel Thread Ring Gauge | 2.5 mm - 160 mm | 2.5 ?m | 5,000/piece | Contact scanning with measuring probe | 08041 - 11702 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
72 | Go - No Go Parallel Thread Ring Gague | from pitch 0.45 - 8 mm | 1.5 µm | 3,500/piece | Cal. by Universal length measuring Machine and T-Ball Sttylus | 08041 - 11801 | Dimensional Metrology Department | 3500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
73 | Taper Thread Ring Gauge | from pitch 0.45 - 8 mm | 2.0 µm | 3,500/piece | Cal. by Universal length measuring Machine and T-Ball Sttylus | 08041 - 11901 | Dimensional Metrology Department | 3500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
74 | Parallel Thread Plug Gauge | from pitch 0.2 - 6 mm | 1.5 µm | 3,500/piece | Cal. by Universal length measuring Machine and 3-Wires | 08041 - 12101 | Dimensional Metrology Department | 3500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
75 | Parallel Thread Plug Gauge | 1 mm - 150 mm | 2.5 ?m | 5,000/piece | Contact scanning with measuring probe | 08041 - 12102 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
76 | Go - No Go Parallel Thread Plug Gauge | from pitch 0.2 - 6 mm | 1.5 µm | 3,500/piece | Cal. by Universal length measuring Machine and 3-Wires | 08041 - 12201 | Dimensional Metrology Department | 3500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
77 | Taper Thread Plug Gauge | from pitch 0.45 - 8 mm | 2.0 µm | 3,500/piece | Cal. by Universal length measuring Machine and T-Ball Sttylus | 08041 - 12301 | Dimensional Metrology Department | 3500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
78 | Taper Thread Plug Gauge | 1 mm - 150 mm | 2.5 ?m | 5,000/piece | Contact scanning with measuring probe | 08041 - 12302 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
79 | Taper Thread Ring Gauge | 2.5 mm - 160 mm | 2.5 µm | 5,000/piece | Contact scanning with measuring probe | 08041 - 12303 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
80 | Pitch | 0.1 mm - 20 mm | 1 µm | 5,000/piece | Contact scanning with measuring probe | 08041 - 12304 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
81 | Thread Angle | 7° - 180° | 0.12° | 5,000/piece | Contact scanning with measuring probe | 08041 - 12305 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
82 | Taper | 0 - 0.5 | 0.0024 | 5,000/piece | Contact scanning with measuring probe | 08041 - 12306 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
83 | Pitch of thread gauge | 0 mm to 120 mm | 2.3 µm | 2,030/piece | Cal. By Contour Measuring Machine | 08041 - 12501 | Dimensional Metrology Department | 2030 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
84 | X-Axis Contour Measuring Machine | 0 mm to 120 mm | 1.3 µm | 10,030/piece | Compare with standard glass scale | 08041 - 12502 | Dimensional Metrology Department | 10030 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
85 | Included angle of thread gauge | 0-90 degree | 0.06 degree | 2,030/piece | Cal. By Contour Measuring Machine | 08041 - 12601 | Dimensional Metrology Department | 2030 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
86 | Contour specimen - x-axis (pitch) | up to 200 mm | 2.3 µm | 5,000/5 points + 1,000/point | Direct measurement by using Contour Measuring Machine | 08041 - 12701 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
87 | Contour specimen - z-axis (step, height) | up to 20 mm | 1.6 µm | 5,000/5 points + 1,000/point | Direct measurement by using Contour Measuring Machine | 08041 - 12702 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
201 | Height Master (analog/digital) | 0 to 300 mm | 0.3 mm + 3E-6 L | 13,500/piece | Cal. by Gauge block or Length Bar | 08051 - 10101 | Dimensional Metrology Department | 13500 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
202 | Height Master (analog/digital) | 0 to 600 mm | 0.3 mm + 3E-6 L | 22,500/piece | Cal. by Gauge block or Length Bar | 08051 - 10102 | Dimensional Metrology Department | 22500 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
88 | Check Master/ Step Gauge | 0 to 300 mm | Q[0.20 µm, 7.7E-06 L] | 10,660/piece | Cal. By Laser Interferometer | 08051 - 10201 | Dimensional Metrology Department | 10660 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
89 | Check Master/ Step Gauge | 0 to 600 mm | Q[0.20 µm, 7.7E-06 L] | 13,160/piece | Cal. By Laser Interferometer | 08051 - 10202 | Dimensional Metrology Department | 13160 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
90 | Check Master/ Step Gauge | 0 to 1020 mm | Q[0.20 µm, 7.7E-06 L] | 15,660/piece | Cal. By Laser Interferometer | 08051 - 10203 | Dimensional Metrology Department | 15660 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
91 | Gear standard Spur gear : Profile slope deviation | 25 to 400 mm | 1.0 to 3.0 µm | 7,880/piece | Cal. by Coordinate Measuring Machine | 08051 - 10301 | Dimensional Metrology Department | 7880 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
92 | Gear standard Spur gear : Profile slope deviation | up to 45o | 1.0 to 2.0 µm | 7,880/piece | Cal. by Coordinate Measuring Machine | 08051 - 10401 | Dimensional Metrology Department | 7880 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
93 | Spur gear : Pitch | up to 400 mm | 2.0 µm | 7,880/piece | Cal. by Coordinate Measuring Machine | 08051 - 10501 | Dimensional Metrology Department | 7880 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
94 | Bevel Gear : form deviation | up to 300 mm | 2.0 µm | 7,880/piece | Cal. by Coordinate Measuring Machine | 08051 - 10601 | Dimensional Metrology Department | 7880 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
95 | Ball Plate | up to 1000 mm | Q[0.26 µm, 1.9E-06 L] | 36,000/piece | Cal. by Coordinate Measuring Machine with Laser Interferometer | 08051 - 10701 | Dimensional Metrology Department | 36000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
96 | 2-D Coordinate (Hole Plate) | Up to 1000 mm | Q[0.25 µm, 0.42E-6L] | 106,250/piece | Swing-round method and laser interferometer integrated with CMM | 08051 - 10702 | Dimensional Metrology Department | 106250 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
97 | Depth Micro-Checker | 0 to 150 mm | Q[0.68 µm, 1.1E-06 L] | 9,000/piece | Cal. by Coordinate Measuring Machine | 08051 - 10801 | Dimensional Metrology Department | 9000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
98 | Depth Micro-Checker | 0 to 300 mm | Q[0.68 µm, 1.1E-06 L] | 11,250/piece | Cal. by Coordinate Measuring Machine | 08051 - 10802 | Dimensional Metrology Department | 11250 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
99 | Depth Micro-Checker | 0 to 150 mm | Q[0.48 µm, 0.67E-06 L] | 10,625/piece | Cal. by substitution method using CMM | 08051 - 10803 | Dimensional Metrology Department | 10625 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
100 | Inside Micro-Checker | 0 to 300 mm | Q[0.20 µm, 7.7E-07 L] | 10,660/piece | Cal. by Laser Interferometer | 08051 - 10901 | Dimensional Metrology Department | 10660 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
101 | Inside Micro-Checker | 0 to 600 mm | Q[0.20 µm, 7.7E-07 L] | 13,160/piece | Cal. by Laser Interferometer | 08051 - 10902 | Dimensional Metrology Department | 13160 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
102 | Inside Micro-Checker | 0 to 1020 mm | Q[0.20 µm, 7.7E-07 L] | 15,660/piece | Cal. by Laser Interferometer | 08051 - 10903 | Dimensional Metrology Department | 15660 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
103 | Caliper Checker/ Vernier Checker | 0 to 300 mm | Q[0.20 µm, 7.7E-06 L] | 10,660/piece | Cal. by Laser Interferometer | 08051 - 11001 | Dimensional Metrology Department | 10660 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
104 | Caliper Checker/ Vernier Checker | 0 to 600 mm | Q[0.20 µm, 7.7E-06 L] | 13,160/piece | Cal. by Laser Interferometer | 08051 - 11002 | Dimensional Metrology Department | 13160 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
105 | Caliper Checker/ Vernier Checker | 0 to 1020 mm | Q[0.20 µm, 7.7E-06 L] | 15,660/piece | Cal. by Laser Interferometer | 08051 - 11003 | Dimensional Metrology Department | 15660 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
106 | 3-D Artifacts | Up to X×Y×Z (1200×1000×700) mm | Q[0.70 µm, 2.30E-6 L] | 2,500 + 900/parameter or depend on drawing | Direct measurement by CMM | 08051 - 11011 | Dimensional Metrology Department | 2500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
231 | NIMT Reference Wall | Up to 8.5 m | Q[3.52 um, 0.22E-6 L] | 112,550/Set | Interferometer | 08051 - 11021 | Dimensional Metrology Department | 112550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
232 | Calibration of Standard of Arm CMM (Sphere Ball Beam) | Up to 2.0 m | Q[0.98 um, 2.01E-06L] | 27,550/piece | Re-position measurement by CMM | 08051 - 11031 | Dimensional Metrology Department | 27550 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
233 | Articulated Arm Coordinate Measuring Machines, (AACMM) using tactile probes verification | Up to 2.5 m | 0.06 um, for probing test Q[1.00 um, 2.32E-06L] for length measurement erro | 75,050/piece | ISO10360-12 | 08051 - 11041 | Dimensional Metrology Department | 75050 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2501 | Laser Tracker | Up to 30 m | Q[3.9 ?m, 1.0E-06L] | 87,500 | ISO 10360-10: 2021 | 08051-11051 | Dimensional Metrology Department | 87500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
203 | Optical Parallel : Flatness | max. dia. 60 mm | 0.02 um | 1,600/piece | Cal. by Flatness Interferometer, Universal Length Measuring Machine (ULM) and/or Gauge Block Comparator (GBC) | 08061 - 10101 | Dimensional Metrology Department | 1600 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
204 | Optical Parallel : Parallelism | max. dia. 60 mm | ULM : Q[0.060 um, 6.48E-06 P] GBC : 0.080 um | 1,600/piece | Cal. by Flatness Interferometer, Universal Length Measuring Machine (ULM) and/or Gauge Block Comparator (GBC) where P representing measured parallelism of the optical parallel | 08061 - 10101 | Dimensional Metrology Department | 1600 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
205 | Optical Parallel : Thickness | max. dia. 60 mm | 0.15 um | 1,600/piece | Cal. by Flatness Interferometer, Universal Length Measuring Machine (ULM) and/or Gauge Block Comparator (GBC) where L representing nominal thickness of the optical parallel | 08061 - 10101 | Dimensional Metrology Department | 1600 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
107 | Optical Flat | max. dia. 60 mm | 20 nm | 2,500/piece | Cal. by Flatness Interferometer | 08061 - 10201 | Dimensional Metrology Department | 2500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
108 | Optical Flat | max. dia. 300 mm | Q[37 nm, 2E-09 F] | 7,500/piece | Cal. by Non-contact measurement using Flatness Interferometer | 08061 - 10202 | Dimensional Metrology Department | 7500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
109 | Transmission Flat/Reference Flat | max. dia. 300 mm | Q[25 nm, 2.0E-06 F] | 11,500/piece | Absolute Flatness calibration by 3-Flat Test method | 08061 - 10203 | Dimensional Metrology Department | 11500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
110 | Glass Hemisphere Roundness specimen (High precision) | dia. Up to 355 mm | Q[8.0 nm, 8.0E-03 R] | 11,500/piece | Error seperation technique by multistep method (Stylus-on-spindle) where R representing roundness value Soft file of measurement result charge +2500 baht/file | 08061 - 10301 | Dimensional Metrology Department | 11500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
111 | Roundness of specimen/workpiece | dia. Up to 355 mm | Q[11.1 nm, 1.1E-02 R] | 7,530/piece | Soft file of measurement result charge +2500 baht/file where R representing measured value of roundness | 08061 - 10401 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
112 | Cylinder/Cylindrical standard - Cylindricity | Height up to 500 mm Diameter up to 300 mm Maximum runout: 2.06 mm | Q[0.16 µm, 1.17E-06 H] | 7,530 | Direct measurement by Roundness Measuring Instrument where H representing measured height of the cylinder | 08061 - 10402 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
113 | Cylinder/Cylindrical standard - Straightness | Height up to 500 mm Diameter up to 300 mm Maximum runout: 2.06 mm | Q[77 nm, 0.10E-06 L] | 7,530 | Direct measurement by Roundness Measuring Instrument based on Reversal method where L representing travel length | 08061 - 10403 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
114 | Cylinder/Cylindrical standard - Parallelism | Height up to 500 mm Diameter up to 300 mm Maximum runout: 2.06 mm | Q[0.16 µm, 1.2E-06 H] | 7,530 | Direct measurement by Roundness Measuring Instrument where H representing measured height of the cylinder | 08061 - 10404 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
115 | Cylinder/Cylindrical standard - Squareness | Height up to 500 mm Diameter up to 300 mm Maximum runout: 2.06 mm | Q[0.40 µm, 1.2E-06 H] | 7,530 | Direct measurement by Roundness Measuring Instrument where H representing measured height of the cylindrical square | 08061 - 10405 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
116 | Harmonic Standard | Up to 80 mm (Diameter up to 300 mm, Height up to 500 mm) | Q[53 nm, 8.0E-3 R] | 7,530/range | Direct measurement by roundness measuring instrument (TR595H machine) | 08061 - 10406 | Dimensional Metrology Department | 7530 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
117 | Roughness Specimen (Type A), Step Height | 25 nm to 32 um | Q[6.7 nm, 5.6E-03 Zm] | 5,030/step | Contact measurement by Roughness (stylus) Measuring Instrument where d representing depth | 08061 - 10501 | Dimensional Metrology Department | 5030 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
118 | Roughness Specimen (Type A), Step Height | 10 nm to 0.1 um | Q[6.0 nm, 0.20 Zm] | 6,280/step | Non-contact measurement using Interference Microscope where d representing depth | 08061 - 10502 | Dimensional Metrology Department | 6280 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
119 | Roughness Specimen (Type C or Type D), Roughness Specimen | 25 nm to 32 um | Ra: Q[9 nm, 5.6E-03 Zm] | 7,530/range | Contact measurement by Roughness (stylus) Measuring Instrument where Ra, Rz and Sm representing roughness value | 08061 - 10601 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
120 | Roughness Specimen (Type C or Type D), Roughness Specimen | 25 nm to 32 um | Rz: Q[21 nm, 8.8 E-03 Zm] | 7,530/range | Contact measurement by Roughness (stylus) Measuring Instrument where Ra, Rz and Sm representing roughness value | 08061 - 10602 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
121 | Roughness Specimen (Type C or Type D), Roughness Specimen | 25 nm to 32 um | Rsm: Q[0.58 um, 20E-03 Sm] | 7,530/range | Contact measurement by Roughness (stylus) Measuring Instrument where Ra, Rz and Sm representing roughness value | 08061 - 10603 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
123 | Roughness specimen (Areal parameter) | 0.01 um to 3 um | Sa: Q[30 nm, 7.3E-03 Zm] | 7,530/range | Non-contact measurement using Interference Microscope where Zm representing measured value | 08061 - 10611 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
124 | Roughness specimen (Areal parameter) | 0.01 um to 3 um | Sq: Q[31 nm, 7.3E-03 Zm] | 7,530/range | Non-contact measurement using Interference Microscope where Zm representing measured value | 08061 - 10612 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
125 | Roughness specimen (Areal parameter) | 0.01 um to 7 um | Sa: Q[52 nm, 7.3E-3 Zm] | 7,530/range | Non-contact and Direct measurement using 3D Measuring Laser Microscope where Zm representing measured value | 08061 - 10621 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
126 | Roughness specimen (Areal parameter) | 0.01 um to 7 um | Sz: Q[166 nm, 1.1E-2 Zm] | 7,530/range | Non-contact measurement using 3D Measuring Laser Microscope where Zm representing measured value | 08061 - 10622 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
236 | Flatness Tester |  | ³ 40 nm | 200,000/ea. | The flatness tester system, reference flat, software, including training | 08061 - 10701 | Dimensional Metrology Department | 200000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
127 | Step height / depth standard | 1 nm to 100 nm | Q[0.73 nm, 1.8E-06 Zm] | 11,500/piece | Measurement by Metrological Nanomeasuring and Nanopositioning Machine | 08061 - 10901 | Dimensional Metrology Department | 11500 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
128 | Pitch standard | 0.1 um to 100 um | Q[3.4 nm, 2.9E-3 Pm] | 11,500/piece | Measurement by Metrological Nanomeasuring and Nanopositioning Machine | 08061 - 11001 | Dimensional Metrology Department | 11500 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
129 | Radius of curvature | up to 600 mm | Q[0.75 um, 2.2E-05 L] | 2,530 | Non-contact measurement using Fizeau interferometer where L representing measured radius of curvature | 08061 - 11101 | Dimensional Metrology Department | 2530 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
130 | Sphericity and form error of spherical part | F/#: 10.7 | 23 nm | 7,530/piece | Non-contact measurement using Fizeau interferometer | 08061 - 11102 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
131 | Lens : Sag | Aperture: 100 mm | 1 um | 2,530/piece | Non-contact measurement using Fizeau interferometer | 08061 - 11103 | Dimensional Metrology Department | 2530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
132 | Lens : Focal length | Concave radius: up to 600 mm Convex radius: up to 600 mm | (7.87 + 0.47E-03 F) um | 3,780/piece | Non-contact measurement using Fizeau Interferometer where F representing calculated focal length Reflective Index (n) of a lens is required | 08061 - 11201 | Dimensional Metrology Department | 3780 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
133 | Lens : Opthalmic power | Concave radius: up to 600 mm Convex radius: up to 600 mm | (0.0006 + 0.0005 D) m-1 | 3,780/piece | Non-contact measurement using Fizeau Interferometer where D representing calculated ophthalmic power Reflective Index (n) of a lens is required | 08061 - 11202 | Dimensional Metrology Department | 3780 | Dimensional Metrology Department | 7 | meter, power, phase, ac, wattmeter
134 | Lens : Prism Diopter | up to 0.3 | Q[0.008, 6.6E-04 ?] | 5,000/piece | Prism diopter calibration using Mechanical-Bearing Rotary Stage Reflective Index (n) of a lens is required | 08061 - 11203 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
135 | Lens ¦Back focal length¦ | 33 mm to 250 mm (Diameter up to 100mm) | Q[0.032 mm, 2.5E-6 F2] | 2,530/piece | Non contact method by Fizeau inteferometer with a sliding | 08061 - 11204 | Dimensional Metrology Department | 2530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
136 | Lens ¦Lens power¦ | 4 m-1 to 30 m-1 (Diameter up to 100mm) | Q[0.003 m-1, 2.7E-5 P2] | 2,530/piece | Non contact method by Fizeau inteferometer with a sliding | 08061 - 11205 | Dimensional Metrology Department | 2530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
137 | Angular deviation of a centerline of cylindrical lens | Up to 30¢ | 1.8¢ | 2,530/piece | Direct measurement with a calibrated CCD camera using a standard glass scale | 08061 - 11206 | Dimensional Metrology Department | 2530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Refractive Index of Spherical Lens | 1.524 to 1.526 | 0.00033 | 7,500 |  | 08061 - 11207 | Dimensional Metrology Department | 7500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
122 | Roughness Specimen (Type C or Type D), Roughness Specimen | 25 nm to 32 um | Rq: Q[11 nm, 4.2E-03 Zm] | 7,530/range | Contact measurement by Roughness (stylus) Measuring Instrument where Zm representing measured value | 08061-10604 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2506 | Non-contact probe | Length up to 20 um for Dynamic mode | 0.2 um + 1.2E-05 L | 1,500/piece | Direct measurement using Dynamic mode, and camparing its reading with laser interferometer | 08061-11301 | Dimensional Metrology Department | 1500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2507 | Non-contact probe | Length up to 4 mm for Static mode | 0.5 um + 4.6E-05 L | 7,530/piece | Direct measurement using Static mode, and camparing its reading with laser interferometer | 08061-11302 | Dimensional Metrology Department | 7530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
207 | Vernier Caliper: Partial surface contact error and Shift error | up to 600 mm | Partial surface contact error: Q[13 um, 1.33E-5 L] Shift error (Internal measuring faces): Q[13 um, 1.33E-5 L] Shift error (Depth and Step measuring faces): Q[13 um, 1.33E-5 L] | 3,380/piece | Cal. by Gauge block and Ring gauge | 08071 - 10201 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
208 | Vernier Caliper: Partial surface contact error and Shift error | 600 mm to 1000 mm | Partial surface contact error: Q[14 um, 1.33E-5 L] Shift error (Internal measuring faces): Q[13 um, 1.33E-5 L] Shift error (Depth and Step measuring faces): Q[13 um, 1.33E-5 L] | 3,380/piece | Cal. by Gauge block and Ring gauge | 08071 - 10202 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
209 | Height Measuring Station | 0 mm to 300 mm | Q[0.70 µm, 1.0E-06 D] | 11,250/piece | Cal. by Gauge Block | 08071 - 10301 | Dimensional Metrology Department | 11250 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
210 | Height Measuring Station | 0 mm to 600 mm | Q[0.70 µm, 1.0E-06 D] | 12,600/piece | Cal. by Gauge Block | 08071 - 10302 | Dimensional Metrology Department | 12600 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
211 | Height Gauge | 0 mm to 600 mm | Q[13 um, 1.3E-5L] | 9,000/piece | Cal. by Gauge block and Electronic comparator | 08071 - 10401 | Dimensional Metrology Department | 9000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
212 | Height Gauge | 0 mm to 1000 mm | Q[14 um, 1.3E-5L] | 12,600/piece | Cal. by Gauge block and Electronic comparator | 08071 - 10402 | Dimensional Metrology Department | 12600 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
213 | External Micrometer | >25 mm to 225 mm | Q[0.6 um, 1.36E-5L] | 2,820/piece | Cal. by Gauge Block | 08071 - 10501 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
214 | External Micrometer | >225 mm to 300 mm | Q[0.7 um, 1.42E-5L] | 3,380/piece | Cal. by Gauge Block | 08071 - 10502 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
215 | External Micrometer (Resolution : < 0.001 mm) | 0 mm to 25 mm | Q[0.1 um, 1.35E-5 L] | 3,380/piece | Cal. by Gauge Block | 08071 - 10504 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
138 | Setting Zero Rod | Up to 1000 mm | (0.22 + 2.3 L2 ) um | 2,030/piece | Cal. by ULM | 08071 - 10601 | Dimensional Metrology Department | 2030 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
216 | Dial gauge (analog/digital) | 0-10 mm | Q[2.6 um, 1.8E-4 L] | 2,820/piece | Cal. by dial gauge tester | 08071 - 10602 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
217 | Dial gauge (analog/digital) | 0-25 mm | Q[2.6 um, 1.8E-4 L] | 3,380/piece | Cal. by dial gauge tester | 08071 - 10603 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
218 | Dial gauge (analog/digital) | 0-50 mm | Q[2.6 um, 1.8E-4 L] | 5,660/piece | Cal. by dial gauge tester | 08071 - 10604 | Dimensional Metrology Department | 5660 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
219 | Dial gauge (analog/digital) | 0-10 mm | Q[0.34 um, 4.5E-5 L] | 2,820/piece | Cal. by ULM | 08071 - 10605 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
220 | Dial gauge (analog/digital) | 0-25 mm | Q[0.34 um, 4.5E-5 L] | 3,380/piece | Cal. by ULM | 08071 - 10606 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
221 | Dial gauge (analog/digital) | 0-50 mm | Q[0.34 um, 4.5E-5 L] | 5,660/piece | Cal. by ULM | 08071 - 10607 | Dimensional Metrology Department | 5660 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
222 | Dial test indicator | up to 1.6 mm | 0.8 µm | 2,820/piece | Cal. by calibration tester | 08071 - 10608 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
223 | Laser displacement sensor | 0-50 mm | Q[12 um, 1.7E-4 L] | 3,380/piece | Cal. by dial gauge tester | 08071 - 10609 | Dimensional Metrology Department | 3380 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
224 | Laser displacement sensor | 0-100 mm | Q[1.2 um, 4E-4 L] | 5,660/piece | Cal. by gauge block | 08071 - 10610 | Dimensional Metrology Department | 5660 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
225 | Dial gauge (analog/digital) | up to 10 mm | Q[0.34 um, 1.8E-4 L] | 2,820/piece | Direct measurement by I-checker | 08071 - 10611 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
226 | Dial gauge (analog/digital) | up to 25 mm | Q[0.34 um, 1.8E-4 L] | 3,380/piece | Direct measurement by I-checker | 08071 - 10612 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
227 | Dial gauge (analog/digital) | up to 50 mm | Q[0.34 um, 1.8E-4 L] | 5,660/piece | Direct measurement by I-checker | 08071 - 10613 | Dimensional Metrology Department | 5660 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
228 | Dial test indicator | up to 1 mm | Q[0.60 ?m, 1.8E-4 L] | 2,820/piece | Direct measurement by I-checker | 08071 - 10614 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
229 | Micrometer Head | 0 mm to 25 mm | Q[0.8 um, 8E-6 L] | 3,380/piece | Direct measurement by length gauge | 08071 - 10701 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
230 | Micrometer Head | 0 mm to 50 mm | Q[0.8 um, 8E-6 L] | 4,500/piece | Direct measurement by length gauge | 08071 - 10702 | Dimensional Metrology Department | 4500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
142 | Calibration Tester | Up to 5 mm | Q[0.13 um, 4.8E-6 L] | 7,880/piece | Cal. by Laser interferometer | 08071 - 10801 | Dimensional Metrology Department | 7880 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
143 | Indicator Checker | up to 100 mm | Q[0.050 ?m, 1.7E-6 L] | 15,000/piece | Cal. by Laser interferometer | 08071 - 10802 | Dimensional Metrology Department | 15000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
139 | Dial Gauge Tester (Scale, Digital) | 0 mm to 25 mm | Q[0.13 um, 8.0E-6 L] | 6,750/piece | Cal. by Standard linear gauge | 08071 - 10901 | Dimensional Metrology Department | 6750 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
140 | Dial Gauge Tester (Scale, Digital) | 0 mm to 60 mm | Q[0.13 um, 8.0E-6 L] | 8,780/piece | Cal. by Standard linear gauge | 08071 - 10902 | Dimensional Metrology Department | 8780 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
141 | Dial Gauge Tester (Scale, Digital) | 0 mm to 100 mm | Q[0.20 um, 8.0E-6 L] | 10,030/piece | Cal. by Standard linear gauge | 08071 - 10903 | Dimensional Metrology Department | 10030 | Dimensional Metrology Department | 5 | digital, dna, analog, plasmid, dial
144 | Electronic Comparator (Mu-Checker / Millitron) (Include: Probe and display unit) Analog M-Checker | 0 mm to 5 mm | 0.2 um | 4,500/piece | Comparison with Laser Hologauge | 08071 - 11201 | Dimensional Metrology Department | 4500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
145 | Electronic comparator: Error of indication | up to 1 mm | Q[0.17 ?m, 1.8E-4 L] | 4,500/piece | Direct measurement by I-checker | 08071 - 11202 | Dimensional Metrology Department | 4500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
146 | Calibration Gauge for Seam Metal | 0 to 5 mm | 0.30 um | 2,530/piece | Cal. By NIMT Line scale interferometer system | 08071 - 11301 | Dimensional Metrology Department | 2530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
147 | Laser Hologauge / Linear gauge | 0 to 10 mm | Q[0.05 um, 1.7E-6L] | 11,250/piece | Cal. by Gauge Block | 08071 - 11401 | Dimensional Metrology Department | 11250 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
148 | Laser Hologauge / Linear gauge | 0 to 50 mm | Q[0.05 um, 1.7E-6L] | 15,750/piece | Cal. by Gauge Block | 08071 - 11402 | Dimensional Metrology Department | 15750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
149 | Linear Gauge: Error of indication | up to 100 mm | Q[0.20 ?m, 8.0E-6 L] | 5,625/piece | Direct measurement by I-checker | 08071 - 11403 | Dimensional Metrology Department | 5625 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
150 | Thickness of specimen/ feeler gauge | up to 10 mm | Q[0.5 um, 9.1E-6 L] | 750/point | Cal. by length gauge | 08071 - 11801 | Dimensional Metrology Department | 750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
151 | Laser Scan Micrometer (Display Unit, Measuring Unit) | Up to 100 mm | 0.5 mm | 11,250/piece | Cal. by Standard plug gauge | 08071 - 11901 | Dimensional Metrology Department | 11250 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
206 | Holtest / 3- point internal micrometer | up to 100 mm | 1.5 mm + 1.5E-6 d | 2,250/piece | Cal. by Ring gauge | 08071 - 12001 | Dimensional Metrology Department | 2250 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
152 | Magnification calibrator / Magnification Checker | 0 mm to 0.4 mm | 0.2 um | 5,630/piece | Direct measurement by linear gauge | 08071 - 12401 | Dimensional Metrology Department | 5630 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
153 | Calibration block | V1 | 0.01 mm | 4,500/piece | Cal. by QV *Depend on the quality of artifact | 08071 - 12601 | Dimensional Metrology Department | 4500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
154 | Calibration block | V2 | 0.01 mm | 3,380/piece | Cal. by QV *Depend on the quality of artifact | 08071 - 12602 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
155 | Calibration block | LSW | 0.01 mm | 3,380/piece | Cal. by QV *Depend on the quality of artifact | 08071 - 12603 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
156 | Scale lupe : Line spacing | 0 - 30 mm | Q[0.0003 mm, 1.7E-6 L] | 2,820/piece + 500 / nex function | Cal. by NIMT Line Scale interferometer system | 08071 - 12701 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
157 | Scale lupe : Diameter/Ellipse | upto 20 mm | Q[0.0005 mm, 3.21E-6 L] | 2,820/piece + 500 / nex function | Cal. By QV | 08071 - 12702 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
158 | Scale lupe : Angle | 0 to 360 degree | 1 minute | 2,820/piece + 500 / nex function | Cal. By QV | 08071 - 12703 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
159 | Objective/Stage micrometer : Line spacing | 0 - 10 mm | Q[0.0003 mm, 1.7E-6 L] | 2,820/piece + 500 / nex function | Cal. by NIMT Line Scale interferometer system | 08071 - 12711 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
160 | Objective/Stage micrometer : Angle | 0 to 360 degree | 1 minute | 2,820/piece + 500 / nex function | Cal. By QV | 08071 - 12712 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
161 | Objective/Stage micrometer : Diameter/Ellipse | upto 5 mm | Q[0.0005 mm, 3.21E-6 L] | 2,820/piece + 500 / nex function | Cal. By QV | 08071 - 12713 | Dimensional Metrology Department | 2820 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
162 | Specimens : Line spacing | 0 - 300 mm | Q[0.0003 mm, 1.7E-6 L] | 3,000 /piece + 500 / nex function | Cal. by NIMT Line Scale interferometer system | 08071 - 12721 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
163 | Specimens : Angle | 0 to 360 degree | 1 minute | 3,000 /piece + 500 / nex function | Cal. By QV | 08071 - 12722 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
164 | Specimens : Diameter/Ellipse | upto 300 mm | Q[0.0005 mm, 3.21E-6 L] | 3,000 /piece + 500 / nex function | Cal. By QV | 08071 - 12723 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
165 | Test sieve | 0.020 - 125 mm | 0.012 - 0.025 mm | 5,000/piece | Cal. By QV | 08071 - 12901 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
188 | Profile projector | Up to 500 mm | Q[0.0015 mm, 0.15E-3 L] | 12,530/piece | Cal by Glass Scale, Angle Gauge Block. | 08071 - 30601 | Dimensional Metrology Department | 12530 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2509 | 2D Scanning Thread Measuring Machine (2D-STMM) | X-axis = 90 mm Z-axis = 180 mm | Q[0.13 ?m, 2.2E-06 L] | 40,000/piece | Direct measurement using Laser Interferometer | 08071-10303 | Dimensional Metrology Department | 40000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
166 | Standard Glass Scale | 0 mm to 300 mm | Q[41 nm, 0.1E-6 L] | 15,750/piece | Cal. by Line Scale Interferometer | 08081 - 10103 | Dimensional Metrology Department | 15750 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
167 | Standard Glass Scale | >300 mm to 500 mm | Q[41 nm, 0.1E-6 L] | 18,000/piece | Cal. by Line Scale Interferometer | 08081 - 10104 | Dimensional Metrology Department | 18000 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
168 | Standard Metal Scale | 0 mm to 1000 mm | 8 mm + 2.3E-6L | 4,500/piece | Cal. By NIMT QV | 08081 - 10201 | Dimensional Metrology Department | 4500 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
169 | Glass scale | 0 mm to 200 mm | Q[0.0006 mm + 1.7E-6 L] | 3,380/piece | Cal. By NIMT Line scale interferometer (Price per 10 measuring points) *Depend on the quality of glass scale | 08081 - 10301 | Dimensional Metrology Department | 3380 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
170 | Glass scale | 0 mm to 400 mm | Q[0.0006 mm + 1.7E-6 L] | 5,630/piece | Cal. By NIMT Line scale interferometer (Price per 10 measuring points) *Depend on the quality of glass scale | 08081 - 10302 | Dimensional Metrology Department | 5630 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
171 | Glass scale | 0 mm to 500 mm | Q[0.0003 mm, 1.0E-6 L] | 6,750/piece | Cal. By QV | 08081 - 10303 | Dimensional Metrology Department | 6750 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
172 | Glass scale | 0 mm to 1000 mm | Q[0.0005 mm, 1.0E-6 L] | 7,880/piece | Cal. By QV | 08081 - 10304 | Dimensional Metrology Department | 7880 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
173 | Standard Glass Scale/ Working Standard Scale/Glass scale | 0 mm to 200 mm | Q[0.11 um, 2.9E-6 L] | 6,650/piece | Cal. by Line Scale Interferometer | 08081 - 10305 | Dimensional Metrology Department | 6650 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
174 | Standard Glass Scale/ Working Standard Scale/Glass scale | >200 mm to 500 mm | Q[0.11 um, 2.9E-6 L] | 10,140/piece | Cal. by Line Scale Interferometer | 08081 - 10306 | Dimensional Metrology Department | 10140 | Dimensional Metrology Department | 4 | standard, solution, mass, set, ph
175 | Calibration Grid | 100x100 mm | 0.0003 mm | 4,500/piece | Cal. by QV *Depend on the quality of artifact | 08081 - 10501 | Dimensional Metrology Department | 4500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
176 | Calibration Grid | 200x200 mm | 0.0003 mm | 5,400/piece | Cal. by QV *Depend on the quality of artifact | 08081 - 10502 | Dimensional Metrology Department | 5400 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
177 | Calibration Grid | 300x300 mm | 0.0005 mm | 6,300/piece | Cal. by QV *Depend on the quality of artifact | 08081 - 10503 | Dimensional Metrology Department | 6300 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
178 | Calibration Grid | > 300 mm | 0.0007 mm | 10,800/piece | Cal. by QV *Depend on the quality of artifact | 08081 - 10504 | Dimensional Metrology Department | 10800 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
179 | Linear Scale | Up to 1000 mm | 7 mm | 5,630/piece | Cal. by laser interferometer | 08081 - 10801 | Dimensional Metrology Department | 5630 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
182 | Measuring Microscope | Up to 500 mm | 1 um | 6,750/piece | Cal. by standard glass scale | 08081 - 10901 | Dimensional Metrology Department | 6750 | Dimensional Metrology Department | 2 | pressure, measuring, instruments, gas, pe
190 | 3D non-contact | z-axis up to 100 mm | Q[0.0012 mm, 4.5E-6 L] | 14,500/piece | Cal by glass scale, angle gauge, standard square and gauge blocks | 08081 - 30803 | Dimensional Metrology Department | 14500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
189 | 3D non-contact | x- and y-axis up to 500 mm | Q[0.00015 mm, 0.16E-6 L] | 14,500/piece | Cal by glass scale, angle gauge, standard square and gauge blocks | 08081 - 30803 | Dimensional Metrology Department | 14500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
181 | 3D non-contact | z-axis up to 100 mm | Q[0.0012 mm, 4.5E-6 L] | 14,500/piece | Cal by glass scale, angle gauge, standard square and gauge blocks | 08081 - 30803 | Dimensional Metrology Department | 14500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
180 | 3D non-contact | x- and y-axis up to 500 mm | Q[0.00015 mm, 0.16E-6 L] | 14,500/piece | Cal by glass scale, angle gauge, standard square and gauge blocks | 08081 - 30803 | Dimensional Metrology Department | 14500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
191 | Measuring Microscope | Up to 500 mm | 1 um | 6,750/piece | Cal by standard glass scale | 08081 - 30901 | Dimensional Metrology Department | 6750 | Dimensional Metrology Department | 2 | pressure, measuring, instruments, gas, pe
192 | Scale Calibrator | Up to 1000 mm | Q[0.006 mm, 15E-06 L] | 5,630/piece | Cal by glass scale *Depend on the quality of artifact | 08081 - 30911 | Dimensional Metrology Department | 5630 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
193 | Scale Calibrator | >1000 to 2000 mm | Q[0.0085 mm, 21E-06 L] | 8,445/piece | Cal by glass scale *Depend on the quality of artifact | 08081 - 30912 | Dimensional Metrology Department | 8445 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
234 | Scale Calibrator | Up to 1000 mm | Q[0.006 mm, 15E-06 L] | 5,630/piece | Cal by glass scale *Depend on the quality of artifact | 08081 - 30913 | Dimensional Metrology Department | 5630 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
235 | Scale Calibrator | >1000 to 2000 mm | Q[0.0085 mm, 21E-06 L] | 8,445/piece | Cal by glass scale *Depend on the quality of artifact | 08081 - 30914 | Dimensional Metrology Department | 8445 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2510 | Steel Tape | up to 1800 mm | Class 1: Q[0.033 mm, 4.4E-06 L] | 3,000/piece | Cal by scale calibrator | 08081-30915 | Dimensional Metrology Department | 3000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
237 | Monodisperse Polystyrene Particle | Diameter 150 nm | 10 mL | 5,000/ea. | TRM-M-9001 | 08091 - 10101 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
238 | Monodisperse Polystyrene Particle | Diameter 100 nm | 10 mL | 5,000/ea. | TRM-M-9002 | 08091 - 10102 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
239 | Monodisperse Polystyrene Particle | Diameter 1.5 um | 10 mL | 5,000/ea. | TRM-M-9003 | 08091 - 10103 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
240 | Monodisperse Polystyrene Particle | Diameter 300 nm | 10 mL | 5,000/ea. | TRM-M-9004 | 08091 - 10104 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
241 | Monodisperse Polystyrene Particle | Diameter 500 nm | 10 mL | 5,000/ea. | TRM-M-9005 | 08091 - 10105 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
242 | Monodisperse Polystyrene Particle | Diameter 1.0 um | 10 mL | 5,000/ea. | TRM-M-9006 | 08091 - 10106 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2502 | Contour Specimen | X-axis = 13.5 mm, 40.0 mm, 65.0 mm and 82.5 mm | 0.003 mm | 20000 | Contact method | 08091-10107 | Dimensional Metrology Department | 20000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2503 | Contour Specimen | Y-axis = 2.2 mm and 3.7 mm | 0.002 mm | 20000 |  | 08091-10108 | Dimensional Metrology Department | 20000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2504 | Contour Specimen | Radius = 6.0 mm and 10 mm | 0.002 mm | 20000 |  | 08091-10109 | Dimensional Metrology Department | 20000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
2505 | Contour Specimen | Angle = 90? and 270? | 0.02? | 20000 |  | 08091-10110 | Dimensional Metrology Department | 20000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
183 | Laser Distance Meter | Up to 20,000 mm | 1 mm | 1,800/piece | Cal by laser interferomter | 08101 - 10001 | Dimensional Metrology Department | 1800 | Dimensional Metrology Department | 7 | meter, power, phase, ac, wattmeter
2508 | Laser Distance Meter | Up to 20,000 mm | 1 mm | 5,000/piece | Compare to the laser interferometer with the standard 40-m Rail System | 08101-10001 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 7 | meter, power, phase, ac, wattmeter
6 | Distance | 40 m | 5 ppm | 35,000/distance (for up to 10 points) + 1,000/additional point | Comparison with the readings from the standard 40-m Rail System | 08101-10002 | Dimensional Metrology Department | 35000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
7 | Electronic Distance Meter (EDM) | 4 Distances: 10m., 20m., 30m. and 40m. | - | 40,000/piece | Verification of EDM using 40-m Rail System | 08101-10003 | Dimensional Metrology Department | 40000 | Dimensional Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Electronic Distance Meter (EDM): Distance | 4 Distances: 10m., 20m., 30m. and 40m. | - | 5,000/piece |  | 08101-10004 | Dimensional Metrology Department | 5000 | Dimensional Metrology Department | 7 | meter, power, phase, ac, wattmeter
6 | DC Current Source | 10 µA, 100 µA | 5.0 µA/A | 9,000 THB first point + 2,500 THB next point | Volt Ampere method | 09010-03111 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
7 | DC Current Source | 1 mA, 10 mA, 100 mA, 1 A, 10 A, 20 A | 5.0 µA/A to 14 µA/A | 9,000 THB first point + 2,500 THB next point | Volt Ampere method | 09010-03121 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
8 | DC Current Source | 30 A, 100 A, 200 A | 14 µA/A to 20 µA/A | 9,000 THB first point + 2,500 THB next point | Volt Ampere method | 09010-03131 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
9 | DC Current Meter | 10 µA, 100 µA | 5.0 µA/A | 9,000 THB first point + 2,500 THB next point | Volt Ampere method | 09010-03211 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
10 | DC Current Meter | 1 mA, 10 mA, 100 mA, 1 A, 10 A, 20 A | 5.0 µA/A to 14 µA/A | 9,000 THB first point + 2,500 THB next point | Volt Ampere method | 09010-03221 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
11 | DC Current Meter | 30 A | 14 µA/A | 9,000 THB first point + 2,500 THB next point | Volt Ampere method | 09010-03231 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
1 | DC Voltage Standard | 1.018 V / 10 V | 0.9 mV/V /0.4 mV/V | 12,500/piece + 6,750/next point |  | 09010-10110 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
2 | Standard Cell | 1 V / 1.018 V | 0.9 mV/V | 12,500/piece |  | 09010-10111 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
3 | DC Voltage Standard | 1 V / 1.018 V / 10 V | 0.03 mV/V/ 0.01 mV/V | 36,000/piece | Compared against JJVS maintained by NIMT (Uncertainty depends on noise of Unit Under Calibration) | 09010-10112 | Electrical Metrology Department | 36000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
2581 | DC Voltmeter | -10 mV to 10 mV -100 mV to 100 mV -1 V to 1 V -10 V to 10 V | to 174 ?V/V 1.8 ?V/V to 18 ?V/V 0.18 ?V/V to 1.8 ?V/V 0.02 ?V/V to 0.20 ?V/V | 15,000/range + 7}500/next range | Direct comparison with conventional Josephson voltage standard (Uncertainty depends on noise of Unit Under Calibration) | 09010-10113 | Electrical Metrology Department | 15000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2582 | DC Voltage Standard | 0.1 V 1 V / 1.018 V 10 V | 0.6 ?V/V 0.03 ?V/V 0.01 ?V/V | 40,000/2 points +8,000/next point | Compared against programmable Josephson voltage standard (Uncertainty depends on noise of Unit Under Calibration) | 09010-10114 | Electrical Metrology Department | 40000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
4 | Picoampmeter | 2 pA to 200 mA | 0.1 mA/A to 0.01 mA/A | 6,750 to 11,250/piece | Used volt-ampere method. | 09010-10320 | Electrical Metrology Department | 6750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
5 | Electrometer | 2 nC to 200 µC | 0.1 mC/C | 6,750 to 11,250/ piece |  | 09010-10321 | Electrical Metrology Department | 6750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
12 | AC/DC Transfer Standard | AC-DC Transfer Difference 2 mV to 20 V Frequency: 10 Hz to 1 MHz AC-DC Transfer Difference > 20 V to 1000 V Frequency: 10 Hz to 100 kHz | 8 mV/V to 0.22 mV/V | 90,000/item |  | 09020-10510 | Electrical Metrology Department | 90000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
13 | Thermal Transfer Standard | AC-DC Transfer Difference 0.5 V to 20 V Frequency: 10 Hz to 1 MHz AC-DC Transfer Difference > 20 V to 1000 V Frequency: 10 Hz to 100 kHz | 8 mV/V to 0.22 mV/V | 67,500/item |  | 09020-10511 | Electrical Metrology Department | 67500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
14 | Thermal Voltage Converter Standard | AC-DC Transfer Difference 0.5 V to 20 V Frequency: 10 Hz to 1 MHz AC-DC Transfer Difference > 20 V to 1000 V Frequency: 10 Hz to 100 kHz | 8 mV/V to 0.22 mV/V | 9,000/item |  | 09020-10512 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
15 | AC Measurement Standard | AC-DC Transfer Difference 10 mV to 20 V Frequency: 10 Hz to 1 MHz AC-DC Transfer Difference > 20 V to 1000 V Frequency: 10 Hz to 100 kHz | 8 mV/V to 0.12 mV/V | 67,500 to 78,750 /item | Excluded WB option. | 09020-10520 | Electrical Metrology Department | 67500 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
16 | AC Measurement Standard | AC Voltage 10 mV to 20 V Frequency: 10 Hz to 1 MHz AC Voltage > 20 V to 1000 V Frequency: 10 Hz to 100 kHz | 11 mV/V to 0.18 mV/V | 78,750 to 85,500 /item | Included WB option. | 09020-10521 | Electrical Metrology Department | 78750 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
17 | Alternating Voltage Measurement Standard | AC Voltage 100 mV to 20 V Frequency: 10 Hz to 1 MHz AC Voltage > 20 V to 1000 V Frequency: 10 Hz to 100 kHz | 11 mV/V to 0.18 mV/V | 67500/item |  | 09020-10522 | Electrical Metrology Department | 67500 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
18 | AC Voltage Current Standard | AC Voltage 100 mV to 1000V Frequency 60 Hz and 400 Hz AC Current 100 mA to 30 A Frequency 60 Hz and 400 Hz | 18 mV/V to 0.58 mV/V 50 mA/A to 0.12 mA/A | 18,130 to 21,700/item |  | 09020-10560 | Electrical Metrology Department | 18130 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
19 | Multifunction Transfer Standard | AC Voltage 100 mV to 20 V Frequency: 10 Hz to 1 MHz AC Voltage > 20 V to 1000 V Frequency: 10 Hz to 100 kHz AC Current 100 mV to 10 A Frequency: 10 Hz to 10 kHz | 11 mV/V to 0.18 mV/V 50 mA/A to 0.12 mA/A | 78750/item |  | 09020-10561 | Electrical Metrology Department | 78750 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
20 | AC Current Shunt | AC-DC Transfer Difference 5 mA to 100 A Frequency: 10 Hz to 100 kHz | 18 mA/A to 0.20 mA/A | 6,750 to 11,250/item |  | 09020-10610 | Electrical Metrology Department | 6750 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
21 | AC-DC current transfer standard | 30 A to 100 A | ? 23 mA/A | 11,250 / range | AC-DC current transfer | 09020-10611 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
22 | Low level ac-dc voltage transfer by micro potentiometer | 2 mV to 200 mV | ? 44 mV/V | 66,250 /item | AC-DC voltage transfer | 09020-10612 | Electrical Metrology Department | 66250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
23 | Wideband calibration source characterize | Wideband output 3.2 V Frequency 10 Hz to 30 MHz | 0.058 mV/V to 2.0 mV/V | 11,250 / item | For Calibrator | 09020-10613 | Electrical Metrology Department | 11250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
24 | Wideband function of AC measurement standard | Flatness error 1 mV to 3.2 V Wideband gain @ 1 kHz Wideband gain error Frequency 10 Hz to 30 MHz | 0.084 mV/V to 2.0 mV/V 0.12 mV/V to 2.8 mV/V 0.15 mV/V to 2.8 mV/V | 11,250/piece | For AC measurement standard | 09020-10614 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
25 | AC/DC transconductance amplifier | DC current: 0 A to 100 A AC current: 500 mA to 100 A Frequency 10 Hz to 100 kHz | 27 mA/A to 87 mA/A | 91,250 / item | For AC and DC output | 09020-10615 | Electrical Metrology Department | 91250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
26 | DC Standard Resistor | 1 kW | 0.04 mW/W | 42,500/item | Compared against QHR standard maintain by NIMT using automatic resistance bridge. | 09030-10210 | Electrical Metrology Department | 42500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
27 | Standard Resistor 1 W | 1 W at 50 mA | ? 0.21 mW/W | 8,625/piece | Used a group of 1 W standard. | 09030-10211 | Electrical Metrology Department | 8625 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
28 | Standard Resistor 10 kW | 10 kW at 0.3 mA | ? 1.4 mW/W | 8,625/piece | Used a group of 10 kW standard. | 09030-10212 | Electrical Metrology Department | 8625 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
29 | Standard Resistor 100 mW | 100 ?? at 200 A | As Specification | 10,000/piece | Used DCC bridge and 400 Amp range extender. | 09030-10213 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
30 | Standard Resistor (Calibrate in oil or in air) | 1 mW to 10 kW | As specification | 6,750/item | Used DCC bridge and 100 Amp range extender. | 09030-10214 | Electrical Metrology Department | 6750 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
34 | Standard Resistor (Calibrate in oil or in air) | 1 TW to 100 TW | >1.25 mW/W to 24.4 mW/W | 7,880/item | Using Dual Source High Resistance Ratio Bridge | 09030-10215 | Electrical Metrology Department | 7880 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
32 | Standard Resistor (Calibrate in oil or in air) | >10 kW to 1GW | £ 500 mW/W | 7,880/item | Using Automatic High Resistance Bridge | 09030-10215 | Electrical Metrology Department | 7880 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
33 | Standard Resistor (Calibrate in oil or in air) | 1 GW to 1 TW | 37 µW/W to 1.25 mW/W | 7,880/item | Using Dual Source High Resistance Ratio Bridge | 09030-10215 | Electrical Metrology Department | 7880 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
35 | Standard Resistor (Calibrate in oil or in air) | 0 W to 1 TW | As specification | 2,250/point test |  | 09030-10216 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
36 | Standard Resistor (Calibrate in oil or in air) | 10 TW at 100 V, 200 V 500 V and 1000 V | 3 mW/W | 2,250/point test |  | 09030-10217 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
39 | High Resistance Meter | Upto 100 TW | As specification | 2,250 / point | Direct | 09030-10220 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
40 | Digital Thermometer (indicator) | up to 10 MW /DCI 0 - 100 mA | £ 5 mW/W /£ 60 mW/W | (22,500 to 29,250)/item |  | 09030-10221 | Electrical Metrology Department | 22500 | Electrical Metrology Department | 5 | digital, dna, analog, plasmid, dial
41 | Digital Thermometer (indicator) | up to 10 MW / DCV 0 -100 mV | ³ 5 mW/W / ³ 20 mW/W | (9,000 to 18,000)/item |  | 09030-10222 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 5 | digital, dna, analog, plasmid, dial
42 | Low Resistance meter | Range 1 kW maximum Current 0.5 A Range 100 W maximum Current 1 A Range 10 W maximum Current 5 A Range 1 W maximum Current 10 A Range 0.1 W maximum Current 20 A Range 0.01 W maximum Current 30 A Range 1 mW maximum Current 100 A | As Specification | 2,250 / point | Direct measurement | 09030-10223 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
43 | High Resistance Meter | Test Voltage: maximum 5000 V Range 10 MW, 20 MW and 50 MW Range 100 MW, 200 MW and 500 MW Range 1 GW, 2 GW and 5 GW Range 10 GW, 20 GW and 50 GW Range 100 GW, 200 GW and 500 GW | As Specification | 2,250 / point |  | 09030-10224 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
44 | Low Resistance meter (High Current) | :100 mW, 200 mW, 400 nW, 600 mW and 800 mW@ Max.120 A : 1 mW, 2 mW, 4 mW, 6 mW and 8 mW @ Max 100 A : 10 mW, 20 mW, 40 mW, 60 mW and 80 mW @ Max 60 A : 0.1 W, 0.2 W, 0.4 W , 0.6 W and 0.8 W @ Max 10 A : 1 W, 2 W, 4 W, 6 W, 8 W and 16 W @ Max 1 A : 40 W, 80 W @ Max 1 A | As Specification | 2,250 / point |  | 09030-10225 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
45 | Low Resistance Meter | Range 1 m? to 10 M? | As Specification | 2,250 / point |  | 09030-10226 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
46 | High Resistance meter | Range 10 M? to 10 T? Maximum test Voltage: 5 kV | As Specification | 2,250 / point |  | 09030-10227 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
47 | Decade Resistance 0.001 W to <1 W | 0.001 W to <1 W | As Specification | 7,000/knob | Direct measurement | 09030-10228 | Electrical Metrology Department | 7000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
48 | Decade Resistance 1 W to <1 MW | 1 W to <1 MW | As Specification | 10,000/knob | Direct measurement | 09030-10229 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
49 | Decade Resistance 1 MW to <10 TW | 1 MW to <10 TW | As Specification | 10,000/knob | Direct measurement | 09030-10230 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
31 | Standard Resistor (Calibrate in oil or in air) | 0.1 mW to 100 mW | As Specification | 7,880/item | Voltage-Ratio Method | 09030-10231 | Electrical Metrology Department | 7880 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
37 | Resistance Calibrator | 0 W to 100 MW | As specification | 20,000/piece |  | 09031-10210 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
38 | Standard Resistor | Test Voltage: maximum 2500 V / Range 100 MW to 1 TW | As Specification | 2,250 / point | Used Volt Amp Method | 09031-10211 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
163 | AC Resistance | 1 W, 25 W, 100 W, 400 W at 60 Hz or 75 Hz 1 kW, 10 kW, 12.9 kW (100 Hz to 10 kHz) | As Specification 2 mW/W - 60 mW/W | 9,000/price/frequecy + 1,250 next frequecy | Direct substitution measurement | 09040-10410 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
164 | AC Resistor | 1 W to 100 kW @ 1000 Hz | > 0.20 mW/W | 5,000/piece | Measured by RLC digibridge and/or precision LCR meter, direct measurment, at desired frequencies. | 09040-10411 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
165 | Additional Measurement at another Frequency | 1 W to 100 kW @ 100 Hz to 100 kHz | > 0.20 mW/W | 1,250 per frequency |  | 09040-10412 | Electrical Metrology Department | 1250 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
166 | Decade AC Resistor Additional Measurement at another Frequency | 1 W to 100 kW @ 1000 Hz | > 0.20 mW/W | 5,000 first decade + 1,250 next step | Measured by RLC digibridge and/or precision LCR meter, direct measurment, at desired frequencies. | 09040-10413 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
167 | Decade AC Resistor Additional Measurement at another Frequency | 1 W to 100 kW @ 100 Hz to 100 kHz | > 0.20 mW/W | (4,500 first decade + 1,250 next step) per one frequency |  | 09040-10414 | Electrical Metrology Department | 4500 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
168 | Fixed, Fused Silica Dielectric Standard Capacitors | 10 pF @ 1000 Hz | 0.087 mF/F | 20,000/piece/frequency | Comparing with a group of 10 pF standards | 09040-10420 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
169 | Fixed, Fused Silica Dielectric Standard Capacitors | 1 pF @ 1000 Hz | 0.71 mF/F | 17,500/piece/ frequency | Applied direct substitution method with the reference standard capacitors. | 09040-10421 | Electrical Metrology Department | 17500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
170 | Fixed, Fused Silica Dielectric Standard Capacitors | 10 pF @ 1000 Hz | 0.70 mF/F | 17,500/piece/ frequency | Applied direct substitution method with the reference standard capacitors. | 09040-10421 | Electrical Metrology Department | 17500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
171 | Fixed, Fused Silica Dielectric Standard Capacitors | 100 pF@ 1000 Hz | 0.70 mF/F | 17,500/piece/ frequency | Applied direct substitution method with the reference standard capacitors. | 09040-10421 | Electrical Metrology Department | 17500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
175 | Fixed Three Terminal, Nitrogen Dielectric Standard Capacitors | 10000 pF @ 1000 Hz | 5.3 mF/F | 20,000/piece/frequency | Comparing with a group of 10 nF standards | 09040-10422 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
174 | Fixed Three Terminal, Nitrogen Dielectric Standard Capacitors | 10000 pF @ 1000 Hz and 1592 Hz | 7.0 mF/F | 12,500/piece/ frequency |  | 09040-10422 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
173 | Fixed Three Terminal, Nitrogen Dielectric Standard Capacitors | 1000 pF @ 1000 Hz and 1592 Hz | 3.5 mF/F | 12,500/piece/ frequency |  | 09040-10422 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
172 | Fixed Three Terminal, Nitrogen Dielectric Standard Capacitors | 10 pF 100 pF @ 1000 Hz and 1592 Hz | 1.2 mF/F | 12,500/piece/ frequency | Applied a null method using the basic ratio bridge. | 09040-10422 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
176 | Two or Three Terminal Mica Dielectric Standard Capacitors | 0.001 mF, 0.01 mF, 0.1 mF and 1 mF @ 1000 Hz | > 0.12 mF/F | 5,630/piece | Applied a null method using the basic ratio bridge. | 09040-10423 | Electrical Metrology Department | 5630 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
177 | Additional Measurement at another Frequency | 0.001 mF, 0.01 mF, 0.1 mF and 1 mF @ 100, 120, 200, 400, 500 Hz, 2 kHz, 5 kHz, 10 kHz | > 0.12 mF/F | 1,250/frequency | Applied a null method using the basic ratio bridge. | 09040-10424 | Electrical Metrology Department | 1250 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
180 | Four-Terminal Pair Capacitance Standard | 1 pF, 10 pF, 100 pF, 1000 pF, 0.01 mF, 0.1 mF and 1 mF @ 1000 Hz | > 0.12 mF/F | 5,630/piece | Measured by precision capacitance bridge, direct measurement, at desired frequency. | 09040-10425 | Electrical Metrology Department | 5630 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
181 | Additional Measurement at another Frequency | 1 pF, 10 pF, 100 pF, 1000 pF, 0.01 mF, 0.1 mF and 1 mF @ 100 Hz to 100 kHz | > 0.20 mF/F | 1,250/frequency | Measured by RLC Digibridge and/or Precision LCR Meter, direct measurement, at desired frequency. | 09040-10426 | Electrical Metrology Department | 1250 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
182 | Fixed Standard Capacitors | 1 pF to 1 F @ 1000 Hz | > 0.20 mF/F | 4,500/piece | Measured by RLC Digibridge and/or Precision LCR Meter, direct measurement, at desired frequency. | 09040-10427 | Electrical Metrology Department | 4500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
183 | Additional Measurement at another Frequency | 1 pF to 1 F @ 12 Hz to 1 MHz | > 0.20 mF/F | 1,250/frequency |  | 09040-10428 | Electrical Metrology Department | 1250 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
184 | Decade capacitor | 1 pF to 10 mF @ 1000 Hz | > 0.12 mF/F | 5,000 first decade + 1,250 next step | Applied a null method using the basic ratio bridge. | 09040-10429 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
189 | Fixed Standard Inductors | 100 mH, 1 mH, 10 mH, 100 mH, 1 H and 10 H @ 100 Hz | > 0.16 mH/H | 12,500/piece | Measured by the direct substitution method with the reference standard inductors. | 09040-10430 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
190 | Additional Measurement at another Frequency | 100 mH, 1 mH, 10 mH, 100 mH, 1 H and 10 H @ 12 Hz to 1 MHz | > 0.20 mH/H | 1,250/frequency | Measured by RLC digibridge and/or precision LCR meter, direct measurement, at desired frequencies. | 09040-10431 | Electrical Metrology Department | 1250 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
191 | Fixed standard Inductor | 1 mH to 10 H @ 1 kHz | > 0.20 mH/H | 5,630/piece | Measured by RLC digibridge and/or precision LCR meter, direct measurment, at desired frequencies. | 09040-10432 | Electrical Metrology Department | 5630 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
192 | Additional Measurement at another Frequency | 1 mH to 10 H @ 12 Hz to 1 MHz | > 0.20 mH/H | 1,250/frequency |  | 09040-10433 | Electrical Metrology Department | 1250 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
193 | Decade Inductors | 1 mH to 10 H @ 1000 Hz | > 0.20 mH/H | 4,500 first decade + 1,250 next step | Measured by RLC digibridge and/or precision LCR meter, direct measurement, at desired frequencies. | 09040-10434 | Electrical Metrology Department | 4500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
194 | Additional Measurement at another Frequency | 1 mH to 10 H @ 12Hz to 1 MHz | > 0.20 mH/H | (4,500 first decade + 1,250 next step) per one frequency |  | 09040-10435 | Electrical Metrology Department | 4500 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
195 | LCR Meter with 4 terminal pair configuration/without 4 terminal pair configuration | 0.001 mF, 0.01 mF, 0.1 mF and 1 mF @ 100, 120, 200, 400, 500 Hz, 1 kHz, 2 kHz, 5 kHz, 10 kHz 1 pF, 10 pF, 100 pF, 1000 pF @ 1000 Hz* 100 mH, 1 mH, 10 mH, 100 mH, 1 H and 10 H @ 100 Hz and 1000 Hz 1 W, 10 W, 100 W, 1000 W @ 1 kHz 1 kW, 10 kW, 12.9 kW (100 Hz to 10 kHz) | > 0.12 mF/F > 0.12 mF/F > 0.16 mH/H > 0.20 mW/W 0.14 mW/W | To be determined | Calibrated by comparing its reading against the standard capacitors, standard inductors and 4TP standard resistors. The correction functions are used to correct additional error due to the test fixture and test leads before measurement.* For 4 terminal pair configuration | 09040-10436 | Electrical Metrology Department | 7880 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
178 | 2 Terminal Pair Capacitance Standard | 1 pF, 10 pF, 100 pF, 1000 pF at 1000 Hz to 10 MHz | > 100 µF/F | 10,000/Item | S-parameter with resonance frequency | 09040-10437 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
179 | 4 Terminal Pair Capacitance Standard | 1 pF, 10 pF, 100 pF, 1000 pF at 1000 Hz to 30 MHz | > 63 µF/F | 12,500/Item | S-parameter with resonance frequency | 09040-10438 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
196 | Fixed Standard Inductors | 10 mH @ 1 kHz | > 75 mH/H | 20,000/piece | Comparison with a group of inductance standards using a LCR meter | 09040-10439 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
185 | Decade capacitor | 1 pF to 1 F @ 1000 Hz | > 0.20 mF/F | 4,500 first decade | Measured by RLC digibridge and/or precision LCR meter, direct measurment, at desired frequencies. | 09041-10420 | Electrical Metrology Department | 4500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
186 | Additional Measurement at another | 1 pF to 1 F @ 12 Hz to 1 MHz | > 0.20 mF/F | (4,500 first decade |  | 09041-10421 | Electrical Metrology Department | 4500 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
187 | Frequency |  |  | + 1,250 next step) per one frequency |  | 09041-10422 | Electrical Metrology Department | 1250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
188 | Precision Capacitance Bridge | 1 pF, 10 pF, 100 pF @ 1000 Hz or 1592 Hz | ³ 0.40 mF/F | 27,500/item | Calibrated by comparing its reading against the reference standard capacitors. The correction functions are used to correct additional error due to the test fixture and test leads before measurement. | 09041-10423 | Electrical Metrology Department | 27500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
50 | Standard Watt Converter | Volt : 600 V Current : 100 A AC Power : 60 kW Frequency : 50 Hz, 60 Hz Power Factor : + 0 to 1 | > 60 mW/VA | 18,000 to 39,500/item |  | 09050-10710 | Electrical Metrology Department | 18000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
51 | AC Wattmeter, Digital Power Meter | Volt : 1000 V Current : 80 A AC Power : 80 kW Frequency : 15 Hz to 450 Hz Power Factor : + 0 to 1 | > 100 mW/VA | 6,750 to 13,500/Channel |  | 09050-10711 | Electrical Metrology Department | 6750 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
52 | Energy Meter single phase/ three phase with pulse | Active eneregy per phase 0.06525 Ws to 4,800,000 Ws Voltage 15 V to 1000 V Current 0.01 A to 80 A Power Factor 1 to 0.087 (Inductive or Capacitive) Frequency 45 Hz to 65 Hz Measurement time 5s to 60 s | ³ 60 mWs/ Vas to ³ 665 mWs/ Vas | single phase 5,000 to 35,000 / channel three phase 6,250 THB to 51,250 THB |  | 09050-10712 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
53 | Electrical Power Source | Power 1.5 W to 60 kW Frequency 50 Hz to 60 Hz Power Factor 1 to 0 (Inductive or Capacitive) Voltage 3 V to 1000 V Current 0.05 A to 100 A Frequency 15 Hz to 6 kHz | Power ³ 70 mW/VA Voltage ³ 30 mV/V Current ³ 30 mA/A Phase ³ 0.005o | 9,000 to 45,000/item |  | 09050-10713 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
54 | DC Power Meter | Power: 1mW to 20 kW Voltage : 100 mV to 1000 V Current : 10 mA to 20 A | ³ 40 mW/ W | 6,750 to 13,500/Channel |  | 09050-10714 | Electrical Metrology Department | 6750 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
55 | Energy Source | Eneregy 41.67 mWh to 3.333 kWh Power: 0.5 W to 20 kW Voltage: 10 V to 1000 V Current: 0.05 A to 20 A Power Factor: 1 Frequency: 50 Hz to 60 Hz Time : 5 minute to 10 minute | > 0.12 % | 5,000 to 15,000 / item |  | 09050-10715 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
56 | Three Phase Watt Meter, Three Phase AC Power Meter, Three Phase Comparator | Power : 4.5 W to 240 kW Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Power Factor : 1 to 0 (Inductive, Capacitive) Frequency : 50 Hz to 60 Hz | 100 µW/VA to 370 µW/VA | 5,000 to 55,500 /Channel |  | 09050-10716 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
57 | Single phase / Three phase energy meter | Energy : (0.0625 Wh to 5000 kWh) / phase Power : (0.75 W to 50 kW) / phase Voltage : 15 V to 1000 V Current : 0.1 A to 50 A Power Factor : 1 to 0.5 (Inductive ,Capacitive) Frequency : 50 Hz to 60 Hz Time : 300 sec to 100 hours | Single phase @ 300 s 515 µWh/Wh to 565 µWh/Wh Three phase @ 300 s 570 µWh/Wh to 665 µWh/Wh | 5,000 to 30,000 /Channel |  | 09050-10717 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
58 | Single Phase / Three Phase Watt Meter | AC Power : 0 W to 172.8 kW (three phase) | 0.12 % for single phase 0.20 % for three phase | 5,000 to 25,000 / channel | High burden Direct measurement | 09050-10718 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
68 | DC Power Meter (20 A to 200 A) | Power: 200 W to 200 kW Voltage : 10 V to 1000 V Current : 20 A to 200 A | 160 µW/W to 325 µW/W | 7,500 THB (2 points) / channel | 7,500 THB for 2 points 2,500 THB/point for additional point | 09050-10719 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
69 | Wattmeter | Power : 0 W to 12 kW | 50 µW/VA to 65 µW/VA | 5,000 THB to 25,000 THB / channel | Comparison measurement | 09050-10720 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
70 | Energy Meter with Pulse (Single phase or Three phase) | Energy: 1.05 mWh to 833.3 Wh per phase | 90 µWh/Wh to 160 µWh/Wh | 5,000 THB to 30,000 THB/ channel | Pulse comparison | 09050-10721 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
71 | Wattmeter AC Power Meter Wattmeter with AC Curent Clamp AC Power Meter with AC Current Clamp (80 A to 120 A)(1 phase or 3 phase) | Power : 0 W to 120 kW per phase Voltage: 15 V to 1000 V Current : 80 A to 120 A PF: 1 to 0 (inductive or Capacitive) Frequency : 50 Hz to 60 Hz | 135 µW/VA to 205 µW/VA 5.6 mW/W to 6.8 mW/W | 5,000 THB to 25,000 THB/ channel |  | 09050-10722 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
72 | Wattmeter with AC Curent Clamp AC Power Meter with AC Current Clamp (500 A to 6000 A)(1 phase or 3 phase) | Power : 3.75 kW to 6000 kW per phase |  | 7,500 THB for 1st point 3,750 THB/point for additional point / channel | Current coil used | 09050-10723 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
73 | Single Phase AC Wattmeter Single Phase AC Power Meter | Reactive Power : 0 var to 50 kvar | 60 µvar/VA to 125 µvar/VA | 5,000 THB to 25,000 THB / channel | Direct measurement | 09050-10724 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
74 | Three Phase AC Wattmeter Three Phase AC Power Meter Three Phase Power Comparator | Reactive Power : 0 var to 150 kvar | 60 µvar/VA to 125 µvar/VA | 5,000 THB to 25,000 THB | Direct measurement | 09050-10725 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
75 | Single Phase Energy Meter with Pulse Output | Reactive Energy: 0.06525 vars to 4800000 vars varh Voltage: 15 V to 1000 V Current : 0.01 A to 80 A QF: 1 to 0.087 (inductive or capacitive) Frequency : 45 Hz to 65 Hz Time : 5 s to 60 s | ³ 60 mvars/ Vas to | 5,000 to 35,000 THB/ channel | Pulse comparison | 09050-10726 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
76 | Three Phase Energy Meter with Pulse Output | Reactive Energy: 0.19575 vars to 14,400,000 vars varh Voltage: 15 V to 1000 V Current : 0.01 A to 80 A QF: 1 to 0.087 (inductive or capacitive) Frequency : 45 Hz to 65 Hz Time : 5 s to 60 s | ³ 60 mvars/ Vas to ³ 665 mvars/ Vas | 6,250 THB to 51,250 THB | Pulse comparison | 09050-10727 | Electrical Metrology Department | 6250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
77 | Single Phase Energy Meter (No Pulse) | Reactive Energy: 0.0625 varh to 16.667 kvarh Power: 0.75 var to 50 kvar Voltage: 15 V to 1000 V Current : 0.1 A to 50 A QF: 1 to 0.5 (inductive or capacitive) Frequency : 50 Hz to 60 Hz Time : 300 s to 1200 s | 200 µvarh/varh to 600 µvarh/varh | 5,000 THB to 15,000 THB / channel | Energy Packet | 09050-10728 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
78 | Three Phase Energy Meter (No Pulse) | Reactive Energy: 0.1875 varh to 50 kvarh Power: 0.75 var to 50 kvar Voltage: 15 V to 1000 V Current : 0.1 A to 50 A QF: 1 to 0.5 (inductive or capacitive) Frequency : 50 Hz to 60 Hz | 200 µvarh/varh to 600 µvarh/varh | 5,000 THB to 15,000 THB | Energy Packet | 09050-10729 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
79 | AC Voltage of Electrical Power Source | AC Voltage 0.1 V to 1000 V Frequency 10 Hz to 1 kHz | 30 µV/V | 5,000 THB to 50,000 THB/ channel | Digital Sampling Watt Meter, DSWM Method | 09050-10730 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
80 | AC Current of Electrical Power Source | AC Current 0.01A to 100 A Frequency 10 Hz to 1 kHz | 40 µA/A | 5,000 THB to 50,000 THB / channel | Digital Sampling Watt Meter, DSWM Method | 09050-10731 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
81 | AC Voltage of AC power meter, or AC wattmeter | Voltage: 5 V to 1000 V Frequency: 16 Hz to 850 Hz | 0.049 mV/V to 0.31 mV/V | 2,500 THB to 25,000 THB / channel | Direct measurement Method | 09050-10732 | Electrical Metrology Department | 2500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
82 | AC Current of AC power meter, or AC wattmeter | Current: 0.01 A to 20 A Frequency: 16 Hz to 850 Hz Current: > 20 A to 80 A Frequency: 40 Hz to 850 Hz | 0.061 mA/A to 0.177 mA/A | 2,500 THB to 25,000 THB / channel | Direct measurement Method | 09050-10733 | Electrical Metrology Department | 2500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
83 | DC Power of Wattmeter | Power : 0.01 mW to 20 kW Voltage: 0.01 V to 1000 V Current : 0.001 mA to 20 A | 0.15 mW/W to 1.2 mW/W | 2,500 THB to 25,000 THB / channel | Direct measurement Method | 09050-10734 | Electrical Metrology Department | 2500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
84 | DC Voltage of Wattmeter | Voltage: 0.01 V to 1000 V | 0.02 mV/V to 0.16 mV/V | 2,500 THB to 25,000 THB / channel | Direct measurement Method | 09050-10735 | Electrical Metrology Department | 2500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
85 | DC Current of Wattmeter | Current: 0.01 A to 20 A | 0.15 mA/A to 1.2 mA/A | 2,500 THB to 25,000 THB/ channel | Direct measurement Method | 09050-10736 | Electrical Metrology Department | 2500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
86 | AC Voltage for VT with Wattmeter | AC Voltage 5 V to 1000 V | 0.051 mV/Vto 0.086 mV/V | 5,000 THB to 15,000 THB / channel | Direct measurement | 09050-10737 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
91 | AC Current for CT with Wattmeter | AC Current : 0.01 A to 120 A Frequency : 16 Hz to 850 Hz | 0.011 % to 0.028 % | 5,000 THB to 15,000 THB / channel | Direct measurement | 09050-10738 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
93 | DC energy meter | DC energy : 0.000833 Wh to 50 kWh Voltage : 1 V to 500 V Current : 0.01 A to 10 A Time : 300 s to 10 hours | 0.42 mWh/Wh to 4.3 mWh/Wh | 5,000 THB to 15,000 THB / channel | Energy Packet | 09050-10739 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
94 | Three phase wattmeter, Three phase power meter,Three Phase Comparator Wattmeter AC power meter Wattmeter | Three phase active power : 0 W to 115.2 kW Voltage : 60 V to 480 V Current :0.005 A to 160 A Frequency: 45 Hz to 65 Hz PF:1, 0.5 (inductiv,capactivie), 0 (inductiv,capactivie), | 54 µW/VA to 204 µW/VA | 5,000 THB to 20,000 THB | Comparison measurement | 09050-10740 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
95 | Single phase wattmeter with frequency output AC power meter Wattmeter | Single phase active power :0.1305 W to 80 kW Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 45 Hz to 65 Hz PF:1 to 0.087, (inductive or capacitive) | 48 mW/VA to 183 mW/VA | 5,000 THB to 17,500 THB/ channel | Direct with frequency measurement | 09050-10741 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
96 | Three phase wattmeter with frequency output AC power meter Wattmeter | Three phase active power: 0.3915 W to 240 kW Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 45 Hz to 65 Hz PF:1 to 0.087, (inductive or capacitive) | 48 mW/VA to 183 mW/VA | 5,000 THB to 17,500 THB | Direct with frequency measurement | 09050-10742 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
97 | Single phase wattmeter with frequency output AC power meter Wattmeter Wattmeter | Single phase reactive power :0.1305 var to 80 kvar Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 45 Hz to 65 Hz QF:1 to 0.087, (inductive or capacitive) | 48 mvar/VA to 183 mvar/VA | 5,000 THB to 17,500 THB | Direct with frequency measurement | 09050-10743 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
98 | Three phase wattmeter with frequency output AC power meter Wattmeter Wattmeter | Three phase reactive power: 0.3915 var to 240 kvar Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 45 Hz to 65 Hz QF:1 to 0.087, (inductive or capacitive) | 48 mvar/VA to 183 mvar/VA | 5,000 THB to 17,500 THB | Direct with frequency measurement | 09050-10744 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
99 | Ac current clamp meter,Ac clamp meter Clamp meter | AC current Ac current : 10 A to 1500 A Frequency: 16 Hz to 300 Hz | 5.045 mA/A | 5,000 THB to 10,000 THB / channel | Indirect measurement | 09050-10745 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
100 | AC power meter, Wattmeter | Power factor (PF, cos j) : 1 to 0 Reactive power factor (QF, sin j) : 1 to 0 Voltage: 15 V to 1000 V Current: 0.1 A to 80 A Frequency: 45 Hz to 65 Hz | 6×10-6 to 61×10-6 | 5,000 THB to 7,500 THB / channel | Direct measurement | 09050-10746 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
101 | AC power meter,Wattmeter | Frequency: 16 Hz to 850 Hz Voltage: 1 V to 1000 V Current: 0.01 A to 80 A | 0.0011 Hz to 0.050 Hz | 5,000 THB to 7,500 THB / channel | Direct measurement | 09050-10747 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
102 | Single phase wattmeter, Single phase power meter,Single phase power comparator AC power meter Wattmeter | Reactive power : 0 var to 100 kvar Voltage : 15 V to 1000 V Current : 80 A to 100 A Frequency: 45 Hz to 65 Hz QF:1 to 0, (inductive or capacitive) | 134 µvar/VA to 213 µvar/VA | 5,000 THB to 12,500 THB / channel | Direct measurement | 09050-10748 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
103 | Three phase wattmeter, Three phase power meter, Three phase power comparator Wattmeter | Reactive power : 0 var to 300 kvar Voltage : 15 V to 1000 V Current : 80 A to 100 A Frequency: 45 Hz to 65 Hz QF:1 to 0, (inductive or capacitive) | 134 µvar/VA to 213 µvar/VA | 5,000 THB to 12,500 THB | Direct measurement | 09050-10749 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
104 | Single phase wattmeter, Single phase power meter,Single Phase Power comparator,Watt converter | Active power : 0 W to 80 kW Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 16 Hz to 400 Hz PF:1 to 0, (inductive or capacitive) | 60 µW/VA to 540 µW/VA | 5,000 THB to 27,500 THB/ channel | Direct measurement | 09050-10750 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
105 | Three phase wattmeter, Three phase power meter,Three Phase Power comparator, | Active power : 0 W to 240 kW Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 16 Hz to 400 Hz PF:1 to 0, (inductive or capacitive) | 60 µW/VA to 540 µW/VA | 5,000 THB to 27,500 THB | Direct measurement | 09050-10751 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
106 | Single phase wattmeter, Single phase power meter,Single Phase Power comparator,Watt converter | Reactive power : 0 var to 80 kvar Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 16 Hz to 400 Hz QF:1 to 0, (inductive or capacitive) | 60 µvar/VA to 540 µvar/VA | 5,000 THB to 27,500 THB / channel | Direct measurement | 09050-10752 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
107 | Three phase wattmeter, Three phase power meter,Three Phase Power comparator, | Reactive power : 0 var to 240 kvar Voltage : 15 V to 1000 V Current : 0.1 A to 80 A Frequency: 16 Hz to 400 Hz PF:1 to 0, (inductive or capacitive) | 60 µvar/VA to 540 µvar/VA | 5,000 THB to 27,500 THB | Direct measurement | 09050-10753 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
108 | Transformer bridge ratio | Phase displacement error -180 deg to 180 deg or -10800 arcmin to 10800 arcmin Voltage: 10 V to 1000 V Frequency: 45 Hz to 65 Hz | 0.35 minute of arc | 10,000 THB to 47,500 THB / item | Direct comparison | 09050-10754 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
109 | Transformer bridge ratio(Phase displacement error for current) | Phase displacement error | 0.42 minute of arc | 11,250 THB to 36,250 THB / item | Direct comparison | 09050-10755 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
110 | Transformer bridge ratio(Voltage Ratio Error) | Ratio error : -0.2 to 0.2 Voltage: 10 V to 1000 V -180 deg to 180 deg Frequency: 45 Hz to 65 Hz | 0.018×10-2 | 10,000 THB to 47,500 THB / item | Direct comparison | 09050-10756 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
111 | Transformer bridge ratio(Current Ratio Error) | Ratio error : -0.2 to 0.2 Current: 0.1 A to 20 A -180 deg to 180 deg Frequency: 45 Hz to 65 Hz | 0.018×10-2 | 11,250 THB to 36,250 THB / item | Direct comparison | 09050-10757 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2583 | Single phase energy meter with pulse output | Active Energy: 15 Ws to 2,880,000 Ws Voltage: 60 V to 480 V Current : 0.05 A to 100 A PF: 1 to 0.5 (inductive) and 1 to 0.8 (capacitive) Frequency : 45 Hz to 65 Hz Time : 10 s to 60 s | 52 ?Ws/Ws | 7,500 THB to 32,500 THB / Channel | Comparison measurement | 09050-10758 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
2584 | Single phase energy meter with pulse output | Voltage: 60 V to 480 V Current : 0.05 A to 100 A QF: 1(inductive) to 0.5 (inductive) and 1(inductive) to 0.8 (capacitive) Frequency : 45 Hz to 65 Hz Time : 10 s to 60 s | 52 ?vars/vars | 7,500 THB to 32,500 THB/Channel | Comparison measurement | 09050-10759 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
2585 | Three phase energy meter with pulse output | Active Energy: 45 Ws to 8,640,000 Ws Voltage: 60 V to 480 V Current : 0.05 A to 100 A PF: 1 to 0.5 (inductive) and 1 to 0.8 (capacitive) Frequency : 45 Hz to 65 Hz Time : 10 s to 60 s | 52 ?Ws/Ws | 20,000 THB to 86,750 THB | Comparison measurement | 09050-10760 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
2586 | Three phase energy meter with pulse output | Reactive Energy: 45 vars to 8,640,000 vars Voltage: 60 V to 480 V Current : 0.05 A to 100 A QF: 1(inductive) to 0.5 (inductive) and 1 (inductive) to 0.8 (capacitive) Frequency : 45 Hz to 65 Hz Time : 10 s to 60 s | 52 ?vars/vars | 20,000 THB to 86,750 THB | Comparison measurement | 09050-10761 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
61 | Phase Meter | Phase Meter Frequency : 15 Hz to 100 kHz Amplitde (rms) : 50 mV to 120 V | > 0.01o | 9,000 to 27,000/item |  | 09050-10920 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
62 | Phase Source | Phase Angle : 0.000o to 999.999o Frequency : 10 Hz to 100 kHz Amplitde (rms) : 50 mV to 120 V | > 0.005 o | 18,000 to 45,000/item |  | 09050-10921 | Electrical Metrology Department | 18000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
63 | Phase Meter Current Shunt | Range : 0.1 A to 50 A Phase Angle : 0.000° to ± 180° Frequency : 16 Hz to 850 Hz | ? 0.002 o | 5,000 (5 pts)'+ 800 / point |  | 09050-10922 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
64 | Power Analyzer, Harmonics and Flicker | Output Voltage : 1020 V Output Current : 80 A Frequency : 20 Hz and up to 1 kHz Harmonics : 100th, up to 6 kHz Flicker : 16 Hz to 9 kHz | Voltage > 100 mV/V Current > 150 mA/A Power > 200 mW/VA Frequency > 50 mHz/Hz | 6,750 to 13,500/channel |  | 09050-10930 | Electrical Metrology Department | 6750 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
65 | Power Ananlyzer, Dips/ Swells | Voltage: 10 V to 1000 V Current Current : 0.05 A to 20A Dip/Swell amplitude : 10 % to 140 % of nominal amplitude Time : 1 ms to 1 minute Frequency : 16 Hz to 850 Hz | ³ 0.07 % | 2,250 to 9,000/ channel |  | 09050-10931 | Electrical Metrology Department | 2250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
66 | Harmonic Current Meter | Current : 1 A to 120 A Fund. Frequency : 50 Hz and 60 Hz Harmonic current : 0.1 % to 10 % of fundamental Harmonics : 1st to 40th, Har. Frequency : 50 Hz to 2.4 kHz | Harmonic < 1 % of Fundamental 0.07 % Harmonic ? 1 % of Fundamental 5.80 % | 5,000 to 27,500/ channel | High burden Direct measurement | 09050-10932 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
67 | Harmonics Power Source | Harmonics Power (0.0025 to 10) % of Fundamental Power Power Factor of Harmonics 1 to 0.5, Inductive or Capacitive Harmonics order : 2nd to 100th Total Power : 20 W to 50 kW Total Voltage : 50 V to 1000 V Total Current : 0.8 A to 50 A Power Factor of Fundamental 1 to 0.5, Inductive or Capacitive Fundamental Frequency : 50 Hz to 60 Hz | 0.32 mW/W to 8.2 mW/W | 10,000 THB/channel (10 points) | 1) Additional fee: 750 THB/ additional point 2) Digital sampling technique | 09050-10933 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
87 | Harmonic Voltage for Power Source | Voltage 1 V to 1000 V Fund. Freq. 45 Hz to 65 Hz Harmonic Amplitude (0.1 % to 25 %) of Fundamental Harmonic order 0, 2nd to 100th Harmonic Frequency DC to 6 kHz | 45 mV/V | 5,000 THB to 50,000 THB/ item | Digital Sampling Technique | 09050-10934 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
88 | Harmonic Current for Power Source | Current 0.01 A to 100 A Fund. Freq. 45 Hz to 65 Hz Harmonic Amplitude (0.1 % to 25 %) of Fundamental Harmonic order 0, 2nd to 100th Harmonic Frequency DC to 6 kHz | 50 mA/A | 5,000 THB to 50,000 THB / channel | Digital Sampling Technique | 09050-10935 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
92 | Flicker impedance of power quality measurement system | Nominal impedance @ 50 Hz (0.24 +j0.15) W for line impedance (016 + j0.10) W for neutral impedance | 0.03% | 5,000 THB to 15,000 THB / channel | Volt-Ampere method | 09050-10936 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
89 | Harmonic source (Voltage harmonic) | Voltage (harmonic) : 0.01 V to 200 V Total rms voltage : 10 V to 1000 V Fundamental frequency : 50 Hz to 60 Hz Harmonic amplitude : 0.1 % to 20 % of fundamental Phase shift of harmonic : 0 degree Harmonic order : o to 100th Harmonic frequency : 0 Hz to 6 kHz | 0.0045 V/V to 3.5 V/V | 6,250 THB to 11,250 THB / channel | Direct measurement | 09050-10937 | Electrical Metrology Department | 6250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
90 | Harmonic source (Current harmonic) | Current (harmonic) : 0.1 mA to 4 A conditions Total rms current: 0.1 A to 20 A Harmonic amplitude : 0.1 % to 20 % of fundamental Phase shift of harmonic : 0 degree Harmonic order : o to 100th Harmonic frequency : 0 Hz to 6 kHz | 0.0037 A/A to 8.8 A/A | 7,500 THB to 14,250 THB / channel | Direct measurement | 09050-10938 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
59 | AC Power Source (On-site) | AC Power : 0.15 W to 76.8 kW | 100 µW/W to 215 µW/W | 5,000 to 28,000 /Channel | Direct measurement | 09050-30710 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
60 | Single phase energy meter,Single phase power meter,Three phase energy meter,Three phase power meter(On-site) | Energy : (0.01875 Wh to 76.8 kWh)/phase | Single/Three phase (? 5 A) | 2,500 to10,000 /Channel | Comparison measurement 1. ? 5 A : single and three phases 2. > 5 A to 160 A : single phase | 09050-30711 | Electrical Metrology Department | 2500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
255 | Energy meter (Three phase) | Active Energy: 11.25 Ws to 259,200,000 Ws Voltage: 25 V to 600 V Current: 0.005 A to 120 A PF:1 to 0.5, (inductive or capacitive) Frequency: 45 Hz to 65 Hz Time : 60 s to 1200 s Temperature 0°C to +46°C Relative humidity 15 % to 85 % | 0.11×10-2 Ws/Ws to 0.50×10-2 Ws/Ws | 5000 (2 points) / item | Pulse comparison during onsite and onload conditions | 09050-30712 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
126 | DC High Voltage Divider | -10 kV to 10 kV | 30 x 10-6 (Ratio) | < 10 kV 4,500 first point '+ 2,250 next point | Compared against DC high voltage reference system. | 09060-10810 | Electrical Metrology Department | 10 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
127 | DC High Voltage Divider | >10 kV to 100 kV | 30 x 10-6 (Ratio) | > 10 kV 5,630 first point '+ 2,250 next point |  | 09060-10811 | Electrical Metrology Department | 10 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
129 | DC High Voltage Source | -10 kV to 100 kV | 30 mV/V | 9,000 to 18,000/item | Compared against DC high voltage reference system. | 09060-10812 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
131 | DC High Voltage Meter | -10 kV to 100 kV | As specification | 6,750 to 18,000/item | Compared against DC high voltage reference system. | 09060-10813 | Electrical Metrology Department | 6750 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
133 | High Voltage Capacitance & Dissipation Factor | Capacitance 10 pF to 1000 pF Voltage:10 kV to 100 kV Frequency : 50 Hz Dissipation Factor: 50 x 10-6 to 1 x 10-3 | ³ 80 mF/F for Capacitance ³ 50 x 10-6 for Dissipation Factor | 12,500/item | Compared against standard capacitor using AC high voltage capacitance and inductance bridge. | 09060-10820 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
134 | High Voltage Capacitance & Dissipation Factor | Capacitance 10 pF to 1000 pF Voltage:10 kV to 200 kV Frequency : 50 Hz Dissipation Factor: 50 x 10-6 to 1 x 10-3 |  | 17,500/item |  | 09060-10821 | Electrical Metrology Department | 17500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
135 | AC High Voltage Divider | 1 kV to 200 kV @ 50 Hz | ³ 0.3 % (Scale Factor) | 11,500 to 27,500/item | Compared against AC high voltage reference system | 09060-10830 | Electrical Metrology Department | 11500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
136 | AC High Voltage Meter | 1 kV to 200 kV @ 50 Hz | As specification | 9,000 to 27,000/item | Compared against AC high voltage reference system. | 09060-10831 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
139 | AC High Voltage Source | 1 kV to 200 kV @ 50 Hz | ³ 0.3 % | 11,500 to 23,000/item |  | 09060-10832 | Electrical Metrology Department | 11500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
140 | Voltage Transformer | 3.3 kV to 33 kV @ 50Hz/60 Hz | Ratio error ³ 0.01 % Phase displacement ³ 0.3 min | 7,500 first tap '+ 1,900 next tap per frequency, burden and power factor | Compared against voltage transformer reference system. 850 THB for additional point | 09060-10833 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
142 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 100 kV | Scale factor ³ 1.1 % | 23,750/ item | Additional fee for time to chopping | 09060-10840 | Electrical Metrology Department | 23750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
143 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 200 kV | Time parameter ³ 1.7 % | 26,250/ Item | (2.0 to 6.0) ms: 5,000 THB/point for | 09060-10840 | Electrical Metrology Department | 26250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
144 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 300 kV |  | 31,250/ item | MS rated < 200 kV, and 10,000 THB/ | 09060-10840 | Electrical Metrology Department | 31250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
145 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 400 kV |  | 33,750/ item | point for MS rated < 500 kV. | 09060-10840 | Electrical Metrology Department | 33750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
146 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 500 kV |  | 36,250/ item |  | 09060-10840 | Electrical Metrology Department | 36250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
155 | Switching Impulse Voltage Measuring System 250/2500 µs | 400 kV |  | 33,750/ item |  | 09060-10841 | Electrical Metrology Department | 33750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
154 | Switching Impulse Voltage Measuring System 250/2500 µs | 300 kV |  | 31,250/ item |  | 09060-10841 | Electrical Metrology Department | 31250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
153 | Switching Impulse Voltage Measuring System 250/2500 µs | 200 kV | Time parameter ³ 1.7 % | 26,250/ Item |  | 09060-10841 | Electrical Metrology Department | 26250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
152 | Switching Impulse Voltage Measuring System 250/2500 µs | 100 kV | Scale factor ³ 0.8 % | 23,750/ item |  | 09060-10841 | Electrical Metrology Department | 23750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
160 | Partial Discharge Calibrator | 1 pC-1000 pC | ³ 1.7 % | 15,000/ item (7 points x 2 polarities) | Additional fee: 1,500 /item for additional (1 point x 2 polarities) | 09060-10850 | Electrical Metrology Department | 15000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
161 | Current Transformer | 5 A to 1500 A @ 50 Hz/60 Hz | Ratio error ³ 0.01 % Phase displacement ³ 0.3 min | 7,500 first tap '+ 1,900 next tap 'per frequency, burden 'and power factor | Compared against current transformer reference system. 850 THB for additional point | 09060-10860 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
162 | Rogowski Coil (current transducer),Rogowski Current Transducer | 10 A to 1500 A @ 50 Hz/60 Hz | ³ 0.5 % | 7500/item '(5 current levels) 'frequency 50 Hz or 60 Hz | Compared against current transformer reference system. Additional fee: 7500/additional 5 current levels | 09060-10861 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
128 | DC High Voltage Divider (On-site),DC High Voltage Measuring System (On-site) | 200 kV | ? 30 x 10-6(Ratio) | 60,000/item '(5 voltage levels) | Additional fee: (1) 7,500/additional voltage level (2) Linearity Test: 10,000 /for 5 voltage levels | 09060-30810 | Electrical Metrology Department | 60000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
130 | DC High Voltage Source (On-site) | 200 kV | ? 170 mV/V | 60,000/item '(5 voltage levels) | Additional fee: (1) 7,500/additional voltage level (2) Linearity Test: 10,000 /for 5 voltage levels | 09060-30811 | Electrical Metrology Department | 60000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
132 | DC High Voltage Meter (On-site) | 200 kV | ? 30 ?V/V | 60,000/item '(5 voltage levels) | Additional fee: (1) 7,500/additional voltage level (2) Linearity Test: 10,000 /for 5 voltage levels | 09060-30812 | Electrical Metrology Department | 60000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
137 | AC High Voltage Meter, AC High Voltage Measuring System | 100 kV | ³ 0.3 % | 67,500/item ( 5 voltage levels) | Additional fee: (1) 7,500/additional voltage level | 09060-30830 | Electrical Metrology Department | 67500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
138 | (Vrms, Vpeak, Vpeak/Ö2 @ 50 Hz) (On-site) | 200 kV | ³ 0.3 % | 82,500/item ( 5 voltage levels) | (2) Linearity Test: 10,000 /for 5 voltage levels | 09060-30831 | Electrical Metrology Department | 82500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
141 | Transformer Loss Measuring System (On-site) | up to 200 kV, up to 5 kA, 50/60 Hz | 0.03 % @ p.f. = 1.0 | 56,250 for 3 power factors (4 points/power factor) +2,500 for each additional point | On-site calibration of voltage transducers, and current transducers of the transformer loss measuring system are a prerequisite for this service. | 09060-30834 | Electrical Metrology Department | 56250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
147 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 100 kV | Scale factor ³ 1.1 % | 71,250/ item (5 voltage levels) | Additional fee: | 09060-30840 | Electrical Metrology Department | 71250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
148 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 200 kV | Time parameter ³ 2.2 % | 78,750/ Item (5 voltage levels) | (1) 15,000/additional voltage level | 09060-30841 | Electrical Metrology Department | 78750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
149 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 300 kV |  | 93,750/ item (5 voltage levels) | (2) Linearity Test: 20,000/for 5 voltage levels | 09060-30842 | Electrical Metrology Department | 93750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
150 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 400 kV |  | 101,250/ item (5 voltage levels) |  | 09060-30843 | Electrical Metrology Department | 101250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
151 | Lightning Impulse Voltage Measuring System 1.2/50 µs | 500 kV |  | 108,750/ item (5 voltage levels) |  | 09060-30844 | Electrical Metrology Department | 108750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
156 | Switching Impulse Voltage Measuring System 250/2500 µs | 100 kV | Scale factor ³ 1.1 % | 71,250/ item (5 voltage levels) | Additional fee: | 09060-30845 | Electrical Metrology Department | 71250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
157 | Switching Impulse Voltage Measuring System 250/2500 µs | 200 kV | Time parameter ³ 2.2 % | 78,750/ Item (5 voltage levels) | (1) 15,000/additional voltage levels | 09060-30846 | Electrical Metrology Department | 78750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
158 | Switching Impulse Voltage Measuring System 250/2500 µs | 300 kV |  | 93,750/ item (5 voltage levels) | (2) Linearity Test: 20,000/for 5 voltage levels | 09060-30847 | Electrical Metrology Department | 93750 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
159 | Switching Impulse Voltage Measuring System 250/2500 µs | 400 kV |  | 101,250/ item (5 voltage levels) |  | 09060-30848 | Electrical Metrology Department | 101250 | Electrical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
245 | DC magnetic flux density / | 0.1 mT < B £ 5 mT | 8 mT/T | 9,000/item | Compared with the standard | 09070-11020 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
246 | DC magnetic flux density / | 0.1 mT/A < B £ 5 mT/A | 5 mT/T | 9,000 first point + next point 630 | Compared with the standard | 09070-11021 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
247 | DC magnetic flux density / | 0.1 mT < B £ 5 mT | 8 mT/T | 9,000 first point +next point 630 | Compared with the standard | 09070-11022 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
248 | #NAME? | 1 MHz, 5 MHz, | £ 1 x 10-13 | 21,600/item | Phase Comparison | 09080-10221 | Electrical Metrology Department | 21600 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
249 | Frequency Counter, Timer Counter | Up to 1 GHz | 1 x 10-12 to 1 x 10-9 | 7,200/item | Direct measurement | 09080-10231 | Electrical Metrology Department | 7200 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
250 | Frequency Meter | 10 kHz to 225 MHz | 3 x 10-13 to 5 x 10-10 | 7,200/item | Direct measurement | 09080-10232 | Electrical Metrology Department | 7200 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
251 | Digital Tachometer | As specification | As specification | 7,000/item | Direct measurement method by the light in LED | 09080-10322 | Electrical Metrology Department | 7000 | Electrical Metrology Department | 5 | digital, dna, analog, plasmid, dial
252 | Stopwatch | As specification | As specification | 2,250/item |  | 09080-10331 | Electrical Metrology Department | 2250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
254 | Standard Frequency Source for Remote Frequncy Standard | 1 MHz, 5 MHz, 10 MHz | £ 1 x 10-13 | 21,600/item | GPS Common view | 09080-30212 | Electrical Metrology Department | 21600 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
227 | Audio Power Amplifer | As specification | As Specification | 3,500/Freq.Point | Voltage Ratio method | 09090-05131 | Electrical Metrology Department | 3500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
219 | Audio Analyzer | As specification | As specification | 17,500 to 27,500 /item | direct measurement | 09090-09321 | Electrical Metrology Department | 17500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
220 | Measuring Receiver | As specification | As specification | 30,000 to 40000 /Item | direct measurement | 09090-09322 | Electrical Metrology Department | 30000 | Dimensional Metrology Department | 2 | pressure, measuring, instruments, gas, pe
221 | Audio Filter | As specification | As specification | 2,500/Freq.point | direct measurement | 09090-09323 | Electrical Metrology Department | 2500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
222 | Eddy current crack detector | As specification | As specification | 15,000 /item | direct measurement | 09090-09324 | Electrical Metrology Department | 15000 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
231 | Onsite- Normalized Site Attenuation (NSA)/ Anechoic Chamber | 30 MHz to 1 GHz | ?1.2 dB | 71,600/item | ref. CISPR 16-1-4 or equivalent | 09090-10311 | Electrical Metrology Department | 71600 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
232 | Onsite- Site VSWR/ Anechoic Chamber | 1 GHz to 18 GHz | ?2.6 dB | 71,600/item | ref. CISPR 16-1-4 or equivalent | 09090-10312 | Electrical Metrology Department | 71600 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
233 | Onsite- Shielding effectivenesss (SE)/ Commercial Chamber | 10 kHz to 18 GHz | ?2.9 dB | 35,800/item | ref. IEEE std 299-2006 | 09090-10313 | Electrical Metrology Department | 35800 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
206 | RF-electric field probe | 100 MHz to 1 GHz 1 GHz to 6 GHz | 0.7 dB 1.4 dB | 45,000/item | Reference standard field/ substitution with standard field probe | 09090-10314 | Electrical Metrology Department | 45000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
207 | Field Uniformity (Onsite- chamber) | 80 MHz to 1 GHz 1 GHz to 6 GHz | 2.4 dB 2.1 dB | 40,000/item | Onsite | 09090-10315 | Electrical Metrology Department | 40000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
200 | Power meter | As specification | As specification | 11,250/item |  | 09090-11111 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
203 | RF Attenuator- Fix Attenuator/Step Attenuator | 10 MHz to 18 GHz, 10 dB to 80 dB | 0.013 dB to 0.22 dB upto frequency range | 3,200/Freq. Point | IF substituition Method (1) | 09090-11120 | Electrical Metrology Department | 3200 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
204 | RF Attenuator- Attenuation in Coaxial line/Passive device | 100 kHz to 18 GHz, 10 dB to 80 dB | 0.027 dB to 0.20 dB upto frequency range | 3,000/Freq. Point | IF substituition Method (2) | 09090-11121 | Electrical Metrology Department | 3000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
205 | Audio Attenuator | 31.5 Hz to 100 kHz | As Specification | 3,500/Freq.Point |  | 09090-11122 | Electrical Metrology Department | 3500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
201 | Feedthrough Mount | 100 kHz to 18 GHz, 1 mW | 0.5 % to 1.3 % | 21,000/item | Using DC substitution Method | 09090-11130 | Electrical Metrology Department | 21000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
202 | Feedthrough Mount | 9 kHz to 26.5 GHz, 1 mW | As Specification | 23,000/item |  | 09090-11131 | Electrical Metrology Department | 23000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
197 | Power Sensor with reflection coefficient | 9 kHz to 18 GHz | 1.0 % to 2.0 % | 20,000/item | Using DC substitution Method | 09090-11150 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
198 | Power Sensor with reflection coefficient | 9 kHz to 26.5 GHz | As Specification | 25,000/item | Using Direct Method | 09090-11151 | Electrical Metrology Department | 25000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
199 | EPROM power sensor (correction C.F.) | As specification | As Specification | 3,000/item |  | 09090-11152 | Electrical Metrology Department | 3000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
208 | Oscilloscope | As specification | As specification | 15,000/item |  | 09090-11160 | Electrical Metrology Department | 15000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
209 | Spectrum Analyzer | As specification | As specification | 50,000/item |  | 09090-11161 | Electrical Metrology Department | 50000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
210 | Network Analyzer | As specification | As specification | 67,500/item |  | 09090-11162 | Electrical Metrology Department | 67500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
211 | Synthesizer Signal Generator /Synthesized Sweeper | Up to 18 GHz | As specification | 8,000 to 22,500/item |  | 09090-11163 | Electrical Metrology Department | 8000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
212 | Function Generator | As specification up to 200 MHz | As specification | 10,000/item |  | 09090-11164 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
213 | RF Signal Generator | As specification | As specification | 22,500/item 33,750/item | Excluded AM, FM and phase modulation. Included AM and FM modulation | 09090-11170 | Electrical Metrology Department | 22500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
224 | Other of RF Microwave Equipment | As specification | As specification | As Cost |  | 09090-11171 | Electrical Metrology Department | 7880 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
215 | RF Voltmeter | 10 MHz to 1 GHz | 0.004 | 11,250/item | Comparison with TVC standard | 09090-11172 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
216 | RF Voltage Source | As specification | As specification | 13,500/item |  | 09090-11173 | Electrical Metrology Department | 13500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
225 | THD Meter | As specification | As specification | 12,500/Item |  | 09090-11175 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
226 | Range Calibrator | As specification | As specification | 10,000/Item |  | 09090-11176 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
214 | Scalar Reflection Coefficient | 40 kHz to 18 GHz | As specification | 15,000/ item | Magnitude | 09090-11210 | Electrical Metrology Department | 15000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
217 | Directional Coupler/ power splitter | 100 kHz to 18 GHz | As Specification | 15,000/ item | Magnitude | 09090-11250 | Electrical Metrology Department | 15000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
2578 | S-parameter / passive device | 9 kHz to 18 GHz | 0.01-0.03 | 5,000-20,000 / item | Direct method by VNA | 09090-11311 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
218 | Directional Coupler/ power splitter | 10 MHz to 18 GHz | As Specification | 18,000/item | Vectors | 09090-11350 | Electrical Metrology Department | 18000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
230 | Amplifier/ Noise figure | 10MHz to 18GHz | 0.21 dB to 0.44 dB | 22,000/Item | Direct measurement | 09090-11431 | Electrical Metrology Department | 22000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
229 | Antenna Factor (AF)/ Horn antenna | 1 GHz to 18 GHz | ?1.6 dB | 45,000/item | Standard Site Method (SSM) | 09090-11511 | Electrical Metrology Department | 45000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2579 | Antenna Factor / Biconical or Log-periodic Antenna | 30 MHz to 1 GHz | 1.6 dB | 45,000 / item | RAM | 09090-11512 | Electrical Metrology Department | 45000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2580 | Antenna Factor / Passive Loop Antenna | 9 kHz to 30 MHz | 1.1 dB | 45,000 / item | Magnetic Field Method | 09090-11513 | Electrical Metrology Department | 45000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
228 | Total Radiated Power (TRP)/ mmWAVE antenna | 24 GHz to 40 GHz | ?1.4 dB | 20,000/item | Indect Far-field (IFF) | 09090-11531 | Electrical Metrology Department | 20000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
223 | Pulse and sine wave generator | Frequency Range : 9 kHz to 1 GHz | As specification | 50,000/item | Direct measurement method with sampling method | 09090-11611 | Electrical Metrology Department | 50000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
234 | Standard Optical Power Meter | 1 mW to 20 mW @ 488 nm,515 nm 50 uW to 20 mW @ 633 nm. 10 uW to 3 mW @ 1310 nm Band 10 uW to 10 mW @ 1550 nm Band | > 0.5 % | 9,380 first point + 3,500 next point | Applied direct comparison method with standard calorimeter. | 09100-12400 | Electrical Metrology Department | 9380 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
235 | Optical Power Meter | 1 mW to 20 mW @ 488nm,515nm 50 uW to 20 mW @ 633 nm 10 uW to 3 mW @ 1310 nm Band 10 uW to 10 mW @ 1550 nm Band 50 uW to 20 mW @ 633 nm 10 uW to 3 mW @ 1310 nm Band 10 uW to 10 mW @ 1550 nm Band | > 0.7 % | 5,500 first point + 1,750 next point | Applied direct comparison method with standard power meter. | 09100-12401 | Electrical Metrology Department | 5500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
236 | Optical Power Meter | 100 mW to 10 W @10600 nm | ? 2 % | 7,500 first point+2,500 next point | Substitutional Method monitor or beam splitter method | 09100-12402 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
237 | Laser Energy Meter | 1mJ to 100 mJ @ 532 nm, 1064 nm | ? 2.4 % | 7,500 first point+2,500 next point | Substitutional Method monitor or beam splitter method | 09100-12403 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
238 | Average pulse power for laser energy meter | 10 mW to 1 W @ 532 nm, 1064 nm | ? 2.4 % | 7,500 first point+2,500 next point | Substitutional Method monitor or beam splitter method | 09100-12404 | Electrical Metrology Department | 7500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
239 | Variable Optical Attenuator | 0 dB to 24 dB @ 1310 nm Band 0 dB to 24 dB @ 1550 nm Band 0 dB to 24 dB @ 1625 nm Band | 0.05 dB | 10,000/item | Follow the standard procedure in IEC 61300-3-14:2014-Error and repeatability of the attenuation setting of a variable optical attenuator | 09100-12408 | Electrical Metrology Department | 10000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
240 | High Power Laser Meter | Power range: 150 W to 1 kW Wavelength: 1080 nm | 0.02 | 12,500 first point + next point 5,000 | Substitutional Method monitor or beam splitter method | 09100-12409 | Electrical Metrology Department | 12500 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
2587 | Single-Mode Fiber Length | 200 m to ? 2000 m @ 1310 nm, 1550 nm | ? 0.062 m | 18,000 first point + 5,000 next point | Follow the standard procedure in IEC 61746-1:2009 Annex A | 09100-12410 | Electrical Metrology Department | 18000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2588 | Single-Mode Fiber Spectral Attenuation | 30 dB @ 1200 nm to 1700 nm | ? 0.012 dB | 18,000 first point + 5,000 next point | Follow the standard procedure in IEC 60793-1-40 Annex A:Cut-back technique | 09100-12411 | Electrical Metrology Department | 18000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2589 | Single-Mode Fiber Attenuation Coefficient Uniformity | 0 dB/km to 10 dB/km @ 1310 nm, 1550 nm | ? 0.002 dB/km | 19,750 first point + 7,500 next point | Follow the standard procedure in IEC 60793-1-40 Annex C:Backscattering technique, IEC TR 62316:Bi-Directional trace, Attenuation uniformity | 09100-12412 | Electrical Metrology Department | 19750 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2590 | Single-Mode Fiber Reflectance | -20 dB to -35 dB @ 1310 nm, 1550 nm | ? 0.42 dB | 18,500 first point + 6,250 next point | Follow the standard procedure in IEC 61746-1:Annex G.3 | 09100-12413 | Electrical Metrology Department | 18500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
241 | CO2 Medical Laser | 10 mW to 30 W @10600 nm | ? 2.4 % | 5,000 first point+2,500 next point | Direct Measurement Method with Reference laser power meter | 09100-32405 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
242 | Nd-YAG Medical Laser | 1 mJ to 100 mJ @532, 1064 nm | ? 2.4 % | 5,000 first point+2,500 next point | Direct Measurement Method with Reference laser energy meter | 09100-32406 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
243 | Nd-YAG Medical Laser | 10 mW to 1 W @532, 1064 nm | ? 2.4 % | 5,000 first point+2,500 next point | Direct Measurement Method with Reference laser energy meter | 09100-32407 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
244 | Fiber Optic Power Meter | 100 ?W to 1 mW @1310 nm Band 100 ?W to 1 mW @1490 nm Band 100 ?W to 1 mW @1550 nm Band 100 ?W to 1 mW @1625 nm Band | ? 0.6 % | 5,000 first point+2,500 next point | Comparison method (Compared with reference standard) | 09100-32408 | Electrical Metrology Department | 5000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
125 | Nano Voltmeter | DC Voltage: 1 mV to 100 mV | As specification | 11,250 to 13,500/item | 2,500 THB for additional adjustment. | 09110-10120 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
112 | Calibrator | DC Voltage : 0 V to 1 kV AC Voltage : 10 mV to 1 kV @ 10 Hz to 1 MHz AC Voltage : 30 V to 1 kV @ 10 Hz to 100 kHz DC Current : 0 A to 20 A AC Current : 100 mA to 20 A @ 10 Hz to 100 kHz Resistance : 0 W to 1 GW | 40 mV/V to 0.6 mV/V 3.03 mV/V to 20 mV/V | 13,500 to 33,750/item | 2,000 to 9,000 THB/item for additional adjustment. | 09110-11310 | Electrical Metrology Department | 13500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
113 | Wideband Option | Up to 30 MHz | As specification | 11,250/option | 3,000 THB/option for additional adjustment. | 09110-11311 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
114 | Power Option | As specification | As specification | 6,750/option | 2,500 THB/option for additional adjustment. | 09110-11312 | Electrical Metrology Department | 6750 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
115 | Scope Option | 300 MHz | As specification | 13,500/option | 2,000 to 6,500 THB/option for additional | 09110-11313 | Electrical Metrology Department | 13500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
116 | Scope Option | 600 MHz to 1 GHz | As specification | 18,000/option | adjustment. | 09110-11314 | Electrical Metrology Department | 18000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
117 | High Current | DC Current : 30 A AC Current : 30 A @ 10 Hz to 10 kHz | 22 mA/A 85 mA/A to 114 mA/A | 11,250/piece |  | 09110-11315 | Electrical Metrology Department | 11250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
118 | Power Quality | AC Voltage -Delta Amplitude Verification, Static Condition 30 V, 300 V @ 50 Hz, 60 Hz -Delta Amplitude Verification, Flicker Condition 230 V@ 50 Hz Composite Harmonic Verification - Voltage 0.12 V to 450 V @ 50 Hz, 60 Hz Type Wave : I, II, III, IV, Square and NRC 7030 - Current 0.11 A to 9.5 A @ 50 Hz, 60 Hz Type Wave : VI, VII, IECA, IECD and NRC 7030 | 3.20 × 10-5 to 4.10 × 10-5 4.13 × 10-5 | 18,000/option |  | 09110-11316 | Electrical Metrology Department | 18000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
119 | Oscilloscope Calibrator | Up to 600 MHz Up to 2.1 GHz | As specification | 40,000/item | 2,000 to 13,500 THB/item for additional adjustment. | 09110-11317 | Electrical Metrology Department | 40000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
120 | Oscilloscope Calibrator | Up to 2.1 GHz | As specification | 45,000/item |  | 09110-11318 | Electrical Metrology Department | 45000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
121 | Amplifier / Precision Power Amplifier | Up to 1kV 10 Hz up to 100 kHz | As specification | 9,000/item | 2,500 THB/function for additional adjustment. | 09110-11319 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
122 | Digital Multimeter | 6 1/2 digit | As specification | 9,000 to 13,500/item | 5,400 THB for additional adjustment. | 09110-11320 | Electrical Metrology Department | 9000 | Electrical Metrology Department | 5 | digital, dna, analog, plasmid, dial
123 | Digital Multimeter | 7 1/2 digit | As specification | 13,500 to 22,500/item | 6,750 THB for aditional calibrate with AC/DC Shunt using direct measurement method | 09110-11321 | Electrical Metrology Department | 13500 | Electrical Metrology Department | 5 | digital, dna, analog, plasmid, dial
124 | Digital Multimeter | 8 1/2 digit | As specification | 22,500 to 27,000/item |  | 09110-11322 | Electrical Metrology Department | 22500 | Electrical Metrology Department | 5 | digital, dna, analog, plasmid, dial
253 | Lidar Speed Meter (Lidar Gun) | 60 km/h,80 km/h, 100 km/h and 120 km/h | 0.58 km/h | 15,000/item | Compare with Reference Speed Generator | 09120-00001 | Electrical Metrology Department | 15000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Mass Standard | 1 kg | 0.038 mg (Class E0) | 18,750/piece | Comparison in air | 10011 - 10001 | Mechanical Metrology Department | 18750 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard set | 1 mg to 10 mg |  | 22,500/set |  | 10011 - 10101 | Mechanical Metrology Department | 22500 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard | 1 mg to 1 kg | Class E1 | 4,500/piece | Mass and Conventional mass | 10011 - 10101 | Mechanical Metrology Department | 4500 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard set | 20 mg to 100 mg |  | 18,000/set |  | 10011 - 10101 | Mechanical Metrology Department | 18000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard set | 200 mg to 1 g |  | 18,000/set |  | 10011 - 10101 | Mechanical Metrology Department | 18000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard set | 200 g to 1 kg |  | 18,000/set |  | 10011 - 10101 | Mechanical Metrology Department | 18000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard set | 2 g to 10 g |  | 18,000/set |  | 10011 - 10101 | Mechanical Metrology Department | 18000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard set | 20 g to 100 g |  | 18,000/set |  | 10011 - 10101 | Mechanical Metrology Department | 18000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard | 2 kg | Class E1 | 5,070/piece | Mass and Conventional mass | 10011 - 10102 | Mechanical Metrology Department | 5070 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard | 5 kg | Class E1 | 5,630/piece | Mass and Conventional mass | 10011 - 10103 | Mechanical Metrology Department | 5630 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard set | 2 kg to 10 kg |  | 21,710/set |  | 10011 - 10104 | Mechanical Metrology Department | 21710 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard | 10 kg | Class E1 | 5,940/piece | Mass and Conventional mass | 10011 - 10104 | Mechanical Metrology Department | 5940 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard | 20 kg | Class E1 | 6,250/piece | Mass and Conventional mass | 10011 - 10105 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 1 mg to 1 kg | Class E2 | 2,050/piece | Conventional mass | 10011 - 10201 | Mechanical Metrology Department | 2050 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 2 kg | Class E2 | 2,750/piece | Conventional mass | 10011 - 10202 | Mechanical Metrology Department | 2750 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 5 kg | Class E2 | 2,750/piece | Conventional mass | 10011 - 10203 | Mechanical Metrology Department | 2750 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 10 kg | Class E2 | 3,130/piece | Conventional mass | 10011 - 10204 | Mechanical Metrology Department | 3130 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 20 kg | Class E2 | 3,940/piece | Conventional mass | 10011 - 10205 | Mechanical Metrology Department | 3940 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 50 kg | Class E2 | 8,000/piece | Conventional mass | 10011 - 10206 | Mechanical Metrology Department | 8000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 1 mg to 1 kg | Class F1 and lower | 1,690/piece | Conventional mass | 10011 - 10301 | Mechanical Metrology Department | 1690 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 2 kg | Class F1 and lower | 2,250/piece | Conventional mass | 10011 - 10302 | Mechanical Metrology Department | 2250 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 5 kg | Class F1 and lower | 2,750/piece | Conventional mass | 10011 - 10303 | Mechanical Metrology Department | 2750 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 10 kg | Class F1 and lower | 2,750/piece | Conventional mass | 10011 - 10304 | Mechanical Metrology Department | 2750 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 20 kg | Class F1 and lower | 3,130/piece | Conventional mass | 10011 - 10305 | Mechanical Metrology Department | 3130 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 50 kg | Class F1 and lower | 6,250/piece | Conventional mass | 10011 - 10306 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 100 kg | Class F1 and lower | 10,170/piece | Conventional mass | 10011 - 10307 | Mechanical Metrology Department | 10170 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 200 kg | Class F1 and lower | 15,000/piece | Conventional mass | 10011 - 10308 | Mechanical Metrology Department | 15000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 500 kg | Class F1 and lower | 15,000/piece | Conventional mass | 10011 - 10401 | Mechanical Metrology Department | 15000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 1000 kg | Class F2 and lower | 16,250/piece | Conventional mass | 10011 - 10402 | Mechanical Metrology Department | 16250 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Mass Standard or Mass Standard set | 2000 kg | Class M1 and lower | 20,000/piece | Conventional mass | 10011 - 10501 | Mechanical Metrology Department | 20000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Volume magnetic susceptibility of weight | 1 g to 20 kg | In the range of 10% to 20% of magnetic polarization | 2,820/piece | By the susceptometer method | 10012 - 10101 | Mechanical Metrology Department | 2820 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Permanent magnetization of weight (including magnetic susceptibility) | 0.005 µT - 2.5 µT /2.5 µT - 120 µT | 0.5 µT / In the range of 10% to 20% of magnetic polarization | 5,000/piece | By the susceptometer method (Nominal weight range 1 g to 20 kg) | 10012 - 10102 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Permanent magnetization of weight | 0 µT - 250 µT | 10 µT + 0.03 µ0M µ0M = Permanent magnetization µT | 1,250/piece | By gaussmeter (Nominal weight range 1 g to 20 kg) | 10012 - 10103 | Mechanical Metrology Department | 1250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Surface roughness (Ra,Rz) of weight | Up to 50 mm | where Zm (unit nm) being measured value of weight | 7,500/piece | Nominal weight range 1 g to 1 kg | 10013 - 10101 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Inter-laboratory comparison on mass measurement; '- Non Automatic Weighing Instruments (NAWI), Electonic balance" | Single range, Resolution 0.001 mg, Maximum capacity 21 g, Number of scale interval 21,000,000 | NIMT.M.M-09 | 5,000 |  | 10014-10101 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Inter-laboratory comparison on mass measurement; | 1000 kg | NIMT.M.M-12 | 20,250 | Conventional mass | 10014-10102 | Mechanical Metrology Department | 20250 | Mechanical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| #NAME? | 200 g , 1 kg and 20 kg | NIMT.M.M-13 | 9,830 | Conventional mass (200 g ,1 kg- F2 Class and 20 kg-M1 Class) | 10014-10103 | Mechanical Metrology Department | 9830 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 100 mg, 100 g, and 1 kg | NIMT.M.M-F1.01/2020 | 5,000 | Conventional mass (100 mg, 100 g, and 1 kg- F1 Class) | 10014-10104 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Non Automatic Weighing Instruments (NAWI), Electonic balance | > 10 kg to 20 kg | 0.64 mg + (1.59 X 10-7) X W | 5200 | Based on the EURAMET calibration guide No.18 version 4.0 (11/2015) | 10015-30101 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Non Automatic Weighing Instruments (NAWI), Electonic balance | > 5g to 100 g | 0.04 mg + (4.62 X 10-7) X W | 5200 | Based on the EURAMET calibration guide No.18 version 4.0 (11/2015) | 10015-30101 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Non Automatic Weighing Instruments (NAWI), Electonic balance | > 100 g to 500 g | 0.04 mg + (4.87 X 10-7) X W | 5200 | Based on the EURAMET calibration guide No.18 version 4.0 (11/2015) | 10015-30101 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Non Automatic Weighing Instruments (NAWI), Electonic balance | > 500 g to 10 kg | 0.45 mg + (2.49 X 10-7) X W | 5200 | Based on the EURAMET calibration guide No.18 version 4.0 (11/2015) | 10015-30101 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Non Automatic Weighing Instruments (NAWI), Electonic balance | > 20 kg to 65 kg | 0.86 mg + (1.32 X 10-7) X W | 5200 | Based on the EURAMET calibration guide No.18 version 4.0 (11/2015) | 10015-30101 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Non Automatic Weighing Instruments (NAWI), Electonic balance | Up to 5 g | 0.03 mg + (7.22 X 10-10) X W | 5200 | Based on the EURAMET calibration guide No.18 version 4.0 (11/2015) | 10015-30101 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Automatic Catchweighing Instruments (ACI) | > 100.0 g to 1000.0 g | 1.05 g | 5,200/3 test loads + 1,250 next test load | 1. Test loads shall be provided by client 2. In-house method based on the EURAMET - Publishable Summary for 14RPT02 AWICal | 10015-30102 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Automatic Catchweighing Instruments (ACI) | (0.2 to 100.0) g | 0.78 g | 5,200/3 test loads + 1,250 next test load | 1. Test loads shall be provided by client 2. In-house method based on the EURAMET - Publishable Summary for 14RPT02 AWICal | 10015-30102 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Automatic Catchweighing Instruments (ACI) | > 1000.0 g to 1500.0 g | 1.12 g | 5,200/3 test loads + 1,250 next test load | 1. Test loads shall be provided by client 2. In-house method based on the EURAMET - Publishable Summary for 14RPT02 AWICal | 10015-30102 | Mechanical Metrology Department | 5200 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Density and volume of weight ( 1 g to 20 kg) | 1 g/2 g/5 g/10 g/20 g/50 g/100 g /200 g /500 g / 1 kg /2 kg /5 kg /10 kg /20 kg | 120 kg m-3, 0.002 cm3 /60 kg m-3, 0.002 cm3 / 20 kg m-3, 0.002 cm3 / 15 kg m-3, 0.002 cm3 / 10 kg m-3, 0.004 cm3 /5.6 kg m-3, 0.004 cm3/ 4.4 kg m-3, 0.004 cm3 /3.8 kg m-3, 0.007 cm3 /2.7 kg m-3, 0.018 cm3/ 2.6 kg m-3, 0.036 cm3 /6.8 kg m-3, 0.21 cm3 /3.3 kg m-3, 0.25 cm3/2.4 kg m-3, 0.4 cm3/2 kg m-3, 0.7 cm3 | 3,920/piece | Any class which no cavity by hydrostatic weighing method | 10021 - 10101 | Mechanical Metrology Department | 3920 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Density and volume of solid artifact | 2700 kg m-3 to 9000 kg m-3 1.25 cm3 to 110 cm3 | 6/V0.8 kg m-3 (0.04123V+0.4592) mm3 V = volume cm3 | 8,750/piece | Hydrostatic weighing method (Traceable to a solid density standard) | 10021 - 10102 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Hydrometer '- Density '- Specific Gravity '- API '- Brix '- Brumé '- Alcoholometer | 600 to 2000 kg m-3 | 3.22x10-5 d + 0.0207 kg m-3 ËÃ×Íà·ÕÂºà·èÒã¹Ë¹èÇÂÍ×è¹ d = density kg m-3 | 7,040/3 points + 1250 next point | Hydrostatic weighing method (Cuckow method) | 10021 - 10301 | Mechanical Metrology Department | 7040 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Hydrometer '- Density '- Specific Gravity '- API '- Brix '- Brumé '- Alcoholometer |  | 0.25 kg m-3 0.00010 sp/gr 0.030 °API 0.30 °Brix 0.030 Bé 0.10 %vol | 4,230/3 points + 1250 next point |  | 10021 - 10302 | Mechanical Metrology Department | 4230 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Density of liquid | 600 to 1500 kg m-3 | 0.1 to 2.0 kg m-3 | 6,750/piece | By hydrostatic weighing method | 10021 - 10401 | Mechanical Metrology Department | 6750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Density of liquid | 700 to 1600 kg m-3 | 0.10 kg m-3 | 1,250/piece | By oscillation-type density meter | 10021 - 10402 | Mechanical Metrology Department | 1250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Density Meter, Benchtop & Portable Oscillation Type | 1000 to 1600 kg m-3 | 0.030 kg m-3 | 22,000/piece | By direct comparison with reference fluids in accordance with ISO 15212 | 10021 - 10501 | Mechanical Metrology Department | 22000 | Mechanical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Density Meter, Benchtop & Portable Oscillation Type | 650 to 1000 kg m-3 | 0.020 kg m-3 | 22,000/piece | By direct comparison with reference fluids in accordance with ISO 15212 | 10021 - 10501 | Mechanical Metrology Department | 22000 | Mechanical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Inter-laboratory comparison on density measurement '- Artifact: Hydrometer | 9 - 81 ?API 640 – 1320 kg m-3 | NIMT.M.D-01/2018 | 5,000 |  | 10021 - 10601 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| CRM - Density of Water : 20 Litres @ 20°C | 998.20 kg m-3 | 1.00 kg m-3 | 500/piece |  | 10021 - 10701 | Mechanical Metrology Department | 500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| CRM - Density of Water : 20 ml @ 20°C | 998.20 kg m-3 | 1.00 kg m-3 | 300/piece |  | 10021 - 10702 | Mechanical Metrology Department | 300 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Dynamic viscosity of liquid | 100 to 15000 mPa·s | 0.1 | 750/sample | By rotational viscometer method | 10022 - 10101 | Mechanical Metrology Department | 750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 1 kN to 110* kN | ISO 376:2011, 0.0020 % | 16,250/direction | Calibrated by deadweight force | 10030 - 10101 | Mechanical Metrology Department | 16250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 0.1 kN to 1.5 kN | ISO 376:2011, 0.0050 % | 13,500/direction | standard machine. | 10030 - 10102 | Mechanical Metrology Department | 13500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 0.2 kN to 5 kN | ISO 376:2011, 0.0050 % | 13,500/direction | Digital indicator shall be provided by | 10030 - 10103 | Mechanical Metrology Department | 13500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 16 kN to 500 kN | ISO 376:2011, 0.0080 % | 23,000/direction | customer. | 10030 - 10104 | Mechanical Metrology Department | 23000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 50 N to 500 N | ISO 376:2011, 0.0040 % | 8,500/direction |  | 10030 - 10105 | Mechanical Metrology Department | 8500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 10 N to 200 N | ISO 376:2011, 0.0040% | 8,500/direction |  | 10030 - 10106 | Mechanical Metrology Department | 8500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 1 N to 50 N | ISO 376:2011, 0.010% | 8,500/direction |  | 10030 - 10107 | Mechanical Metrology Department | 8500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | 20 kN to 100 kN | Accroding to ISO 376:2011 Class 0.5 to 2 | 11,000/direction | Calibrated by force-comparator | 10030 - 10201 | Mechanical Metrology Department | 11000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Force-proving instrument | > 100 kN to 500 kN |  | 15,000/direction | Digital indicator shall be provided by | 10030 - 10202 | Mechanical Metrology Department | 15000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Strain Amplifier(Digital Measurement Amplifier) | ± 2.0 mV/V ± 5.0 mV/V | 2.5 x 10-5 mV/V 3.0 x 10-4 mV/V | 4,200/range | Calibrated by Bridge Calibration Unit | 10030 - 10401 | Mechanical Metrology Department | 4200 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Bridge Calibration Unit | ± 2.5 mV/N | 0.000070 nV/N | 5,000/range |  | 10030 - 10402 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Universal Testing Machine | 5 N to 2 MN (Compression with increasing and decreasing force) | ISO 7500-1 : 2018, 0.35 % |  |  | 10030 - 30201 | Mechanical Metrology Department | 7880 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Universal Testing Machine | 5 N to 500 kN (Tension with increasing force only) | ISO 7500-1 : 2018, 0.17 % | 11,000/force scale | Calibrated by force-measuring | 10030 - 30201 | Mechanical Metrology Department | 11000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Universal Testing Machine | 5 N to 500 kN (Tension with increasing and decreasing force) | ISO 7500-1 : 2018, 0.24% | /mode | device and ascessories shall be | 10030 - 30201 | Mechanical Metrology Department | 7880 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Universal Testing Machine | 5 N to 2 MN (Compression with increasing force only) | ISO 7500-1 : 2018, 0.047 % |  | provided by customer | 10030 - 30201 | Mechanical Metrology Department | 7880 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Static Torque Measuring Devices(Torque Transducer) | 0.1 N·m to 1 N·m 1 N·m to 10 N·m 10 N·m to 1,000 N·m 100 N·m to 5,000 N·m | DIN 51309, 0.015% DIN 51309, 0.01%, 0.03% DIN 51309, 0.01% | 13,100/range |  | 10040 - 10101 | Mechanical Metrology Department | 13100 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Measuring Devices(Torque Transducer) |  |  | 26,250/range |  | 10040 - 10102 | Mechanical Metrology Department | 26250 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Transfer Wrenches | 1 N·m to 10 N·m 10 N·m to 2,000 N·m | DAkkS-DKD-R 3-7, 0.04% DAkkS-DKD-R 3-7, 0.03% | 13,100/range |  | 10040 - 10202 | Mechanical Metrology Department | 13100 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| #NAME? | 1 N·m to 10 N·m | Based on DAkkS-DKD-R 3-8, 0.2% | 6,550/range |  | 10040 - 10302 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N·m to 2,000 N·m | DAkkS-DKD-R 3-8, 0.2% | 6,550/range |  | 10040 - 10303 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices |  | ¡ÁÇ. 1 - 2560 0.2% | 5,250/range |  | 10040 - 10304 | Mechanical Metrology Department | 5250 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| #NAME? | 0.1 N·m to 1 N·m | Based on DAkkS-DKD-R 3-8, 0.3% | 6,550/range |  | 10040 - 10305 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Hand Torque Tools (Type I and Type II) '- Wrench '- Screwdriver | 1 N·m to 1,000 N·m | ISO 6789, 1% | 4,350/range |  | 10040 - 10401 | Mechanical Metrology Department | 4350 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Hand Torque Tools (Type I and Type II) '- Wrench '- Screwdriver | 0.1 N·m to 1 N·m 1 N·m to 10 N·m | ISO 6789, 1.5% ISO 6789, 1% | 4,350/range |  | 10040 - 10402 | Mechanical Metrology Department | 4350 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Torque multiplier | 340 N·m to 2,700 N·m | Based on ISO 6789 : 2003(E), 1.5% | 9,350/range |  | 10040 - 10501 | Mechanical Metrology Department | 9350 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Recalibration of Reference Block | 20 HRA to 95 HRA | 0.40 HRA | 6,550/piece | Reference Block shall be provided | 10040 - 10601 | Mechanical Metrology Department | 6550 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Standard Torque Transducer | ? 2,000 N?m | DIN 51309 Class 0.2, 0.1, 0.05 | 233,500/piece |  | 10040 - 20501 | Mechanical Metrology Department | 233500 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transducer | ? 50 N?m | DIN 51309 Class 0.2, 0.1, 0.05 | 163,000/piece |  | 10040 - 20502 | Mechanical Metrology Department | 163000 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transducer | ? 500 N?m | DIN 51309 Class 0.2, 0.1, 0.05 | 186,400/piece |  | 10040 - 20502 | Mechanical Metrology Department | 186400 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transducer | ? 200 N?m | DIN 51309 Class 0.2, 0.1, 0.05 | 174,250/piece |  | 10040 - 20502 | Mechanical Metrology Department | 174250 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transducer | ? 1,000 N?m | DIN 51309 Class 0.2, 0.1, 0.05 | 196,150/piece |  | 10040 - 20502 | Mechanical Metrology Department | 196150 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transducer | ? 20 N?m | DIN 51309 Class 0.2, 0.1, 0.05 | 158,650/piece |  | 10040 - 20503 | Mechanical Metrology Department | 158650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench | ? 2,000 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 249,450/piece |  | 10040 - 20601 | Mechanical Metrology Department | 249450 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench | ? 500 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 195,600/piece |  | 10040 - 20602 | Mechanical Metrology Department | 195600 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench | ? 50 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 168,850/piece |  | 10040 - 20602 | Mechanical Metrology Department | 168850 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench | ? 1,000 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 206,750/piece |  | 10040 - 20602 | Mechanical Metrology Department | 206750 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench | ? 200 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 181,750/piece |  | 10040 - 20602 | Mechanical Metrology Department | 181750 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench | ? 20 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 163,900/piece |  | 10040 - 20603 | Mechanical Metrology Department | 163900 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| #NAME? | 1 N·m to 10 N·m | Based on DAkkS-DKD-R 3-8, 0.2% | 6,550/range |  | 10040 - 30302 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N·m to 2,000 N·m | DAkkS-DKD-R 3-8, 0.2% | 6,550/range |  | 10040 - 30303 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices |  | ¡ÁÇ. 1 - 2560 0.2% | 5,250/range |  | 10040 - 30304 | Mechanical Metrology Department | 5250 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| #NAME? | 0.1 N·m to 1 N·m | Based on DAkkS-DKD-R 3-8, 0.3% | 6,550/range |  | 10040 - 30305 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1 N·m to 10 N·m | Based on DAkkS-DKD-R 3-8, 1.1% | 6,550/range |  | 10040 - 30306 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 10 N·m to 100 N·m | Based on DAkkS-DKD-R 3-8, 0.75% | 6,550/range |  | 10040 - 30307 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 0 degree to 360 degree | 0.55 degree | 6,550/range |  | 10040 - 30308 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 500 N·m to 5000 N·m | Based on DIN 51309, 1.8 % | 8,000/range |  | 10040 - 30309 | Mechanical Metrology Department | 8000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 40 N·m to 400 N·m | MT 01/2550 | 12,150 |  | 10040 - 30501 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 8 N·m to 40 N·m | MT 01.1/2551 | 12,150 |  | 10040 - 30502 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 10 N·m to 200 N·m | MT 02/2551 | 12,150 |  | 10040 - 30503 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 20 N·m to 330 N·m | MT 03/2553 | 9,750 |  | 10040 - 30504 | Mechanical Metrology Department | 9750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 400 N·m to 1000 N·m | MT 01.2/2563 | 9,750 |  | 10040 - 30505 | Mechanical Metrology Department | 9750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 200 N·m to 1000 N·m | MT 02.1/2564 | 12,150 |  | 10040 - 30506 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 0.1 N·m to 1 N·m | MT 04/2565 | 9,750 |  | 10040 - 30507 | Mechanical Metrology Department | 9750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1 N·m to 10 N·m | MT 04.1/2565 | 9,750 |  | 10040 - 30508 | Mechanical Metrology Department | 9750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1 N·m to 50 N·m | MT 05/2566 | 12,150 |  | 10040 - 30509 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 50 N·m to 200 N·m | MT 05.1/2566 | 12,150 |  | 10040 - 30510 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 200 N·m to 1000 N·m | MT 05.2/2566 | 12,150 |  | 10040 - 30511 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 10 N·m to 200 N·m | MT 06/2566 | 12,150 |  | 10040 - 30512 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 200 N·m to 1000 N·m | MT 06.1/2566 | 12,150 |  | 10040 - 30513 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 20 N·m to 330 N·m | MT 07/2566 | 9,750 |  | 10040 - 30514 | Mechanical Metrology Department | 9750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Adjustment charge of torque device for on-site calibration |  |  | 100% of the calibration cost of each item and each range |  | 10040 - 30601 | Mechanical Metrology Department | 100 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method) - Run-out (Radial run-out and Lateral run-out) | Nominal section width, RN from 95 mm to 315 mm | Radial run-out = 0.02 mm | 5000 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30701 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method) - Run-out (Radial run-out and Lateral run-out) |  | Lateral run-out = 0.02 mm | 5000 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30702 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method) - Load application (Deviation angular from wheel center) | At Load index, LI = 0 to 121 (TLCC = 46 kgf to 1,450 kgf) and At Load index, LI = 122 to 279 (TLCC = 1,500 kgf to 136,000 kgf) | 0.002° | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30703 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method)- Tyre alignment (Camber angle and Slip angle) | At Load index, LI = 0 to 121 (TLCC = 46 kgf to 1,450 kgf) and At Load index, LI = 122 to 279 (TLCC = 1,500 kgf to 136,000 kgf) | Camber angle = 0.035° | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30704 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method)- Tyre alignment (Camber angle and Slip angle) | At Load index, LI = 0 to 121 (TLCC = 46 kgf to 1,450 kgf) and At Load index, LI = 122 to 279 (TLCC = 1,500 kgf to 136,000 kgf) | Slip angle = 0.030° | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30705 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method) - Tyre load | At Load index, LI = 0 to 121 (TLCC = 46 kgf to 1,450 kgf) and At Load index, LI = 122 to 279 (TLCC = 1,500 kgf to 136,000 kgf) | 0.037% for increasing only 0.11% for increasing and decreasing | 10000 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30706 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method) - Spindle force | At Load index, LI = 0 to 121 (TLCC = 46 kgf to 1,450 kgf) and At Load index, LI = 122 to 279 (TLCC = 1,500 kgf to 136,000 kgf) | 0.037% for increasing only 0.11% for increasing and decreasing | 10,000.00 / Direction | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30707 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method)- Surface speed | 0 km/h to 300 km/h | 0.20 km/h | 5,000.00 / Direction | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30708 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Force method) - Drum diameter | 1,700 mm and 2,000 mm | 0.15 mm | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30709 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method) - Run-out (Radial run-out and Lateral run-out) | Nominal section width, RN from 95 mm to 315 mm | Radial run-out = 0.02 mm | 5000 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30801 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method) - Run-out (Radial run-out and Lateral run-out) |  | Lateral run-out = 0.02 mm | 5000 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30802 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method) - Load application (Deviation angular from wheel center) | At Load index, LI = 0 to 121 (TLCC = 46 kgf to 1,450 kgf) and At Load index, LI = 122 to 279 (TLCC = 1,500 kgf to 136,000 kgf) | 0.001° | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30803 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method)- Tyre alignment (Camber angle and Slip angle) |  | Camber angle = 0.035° | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30804 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method)- Tyre alignment (Camber angle and Slip angle) |  | Slip angle = 0.030° | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30805 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method) - Tyre load |  | 0.037% for increasing only 0.11% for increasing and decreasing | 10000 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30806 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method) - Torque input | 0.1 Nžm to 5 kNžm | 0.015% (0.1 Nžm to 1 Nžm) 0.010% (1 Nžm to 5 kNžm) | 20,000.00 / Range | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30807 | Mechanical Metrology Department | 20000 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Rolling resistance testing machine (Torque method) - Surface speed | 0 km/h to 300 km/h | 0.20 km/h | 5,000.00 / Direction | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30808 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rolling resistance testing machine (Torque method) - Drum diameter | 1,700 mm and 2,000 mm | 0.15 mm | 7500 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30809 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Dynamic growth of tyre testing machine - Run-out (Radial run-out and Lateral run-out) | Nominal rim diameter, d from 102 mm to 584 mm | Radial run-out = 0.02 mm | 6250 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30901 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Dynamic growth of tyre testing machine - Tyre speed | 0 km/h to 300 km/h | 0.20 km/h | 6,250 / Direction | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30903 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Dynamic growth of tyre testing machine - Centrifugal radius | 127 mm to 434 mm (254 mm to 868 mm diameter) | 0.11 mm | 9350 | UN Regulation No. 117 and ÁÍ¡. 2721-2560 | 10040 - 30904 | Mechanical Metrology Department | 9350 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor and Connection profile rotatable, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 18,750/range |  | 10040-10310 | Mechanical Metrology Department | 18750 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor and Connection profile rotatable, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-10311 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor rotatable and Connection profile fixed, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-10312 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor rotatable and Connection profile fixed, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 6,550/range |  | 10040-10313 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor fixed and Connection profile rotatable, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 18,750/range |  | 10040-10314 | Mechanical Metrology Department | 18750 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor fixed and Connection profile rotatable, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-10315 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor and Connection profile fixed, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-10316 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration Devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor and Connection profile fixed, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 6,550/range |  | 10040-10317 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| #NAME? | 1 N·m to 1,000 N·m, Sensor and Connection profile rotatable, target uncertainty < 0.5 %, For torque with cross force application | DKD-R 10-8, 0.40 % | 18,750/range |  | 10040-10318 | Mechanical Metrology Department | 18750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1 N?m to 1,000 N?m, Sensor and Connection profile rotatable, target uncertainty ? 0.5 %, For torque with cross force application | DKD-R 10-8, 0.40 % | 12,650/range |  | 10040-10319 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1 N·m to 1,000 N·m, Sensor rotatable and Connection profile fixed, target uncertainty < 0.5 %, For torque with cross force application | DKD-R 10-8, 0.40 % | 12,650/range |  | 10040-10320 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1 N?m to 1,000 N?m, Sensor rotatable and Connection profile fixed, target uncertainty ? 0.5 %, For torque with cross force application | DKD-R 10-8, 0.40 % | 6,550/range |  | 10040-10321 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1 N·m to 10 N·m, For pure torque application | Base on DIN 51309, 1.0 % | 6,550/range |  | 10040-10322 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | (0-360) degree, For angle | 0.40 degree | 6,550/range |  | 10040-10323 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Standard Torque Transfer Wrench Model: TW-MCS01 | ? 20 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 174,400/piece | 1. Temperature compensation on zero signal 2. Temperature compensation on span signal 3. Zero balance 4. Measurement and check signals | 10040-20604 | Mechanical Metrology Department | 174400 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench Model: TW-MCS01 | ? 50 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 180,600/piece | 1. Temperature compensation on zero signal 2. Temperature compensation on span signal 3. Zero balance 4. Measurement and check signals | 10040-20604 | Mechanical Metrology Department | 180600 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench Model: TW-MCS01 | ? 200 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 196,750/piece | 1. Temperature compensation on zero signal 2. Temperature compensation on span signal 3. Zero balance 4. Measurement and check signals | 10040-20604 | Mechanical Metrology Department | 196750 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench Model: TW-MCS01 | ? 500 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 214,050/piece | 1. Temperature compensation on zero signal 2. Temperature compensation on span signal 3. Zero balance 4. Measurement and check signals | 10040-20604 | Mechanical Metrology Department | 214050 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench Model: TW-MCS01 | ? 2,000 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 281,350/piece | Standard Option | 10040-20604 | Mechanical Metrology Department | 281350 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Standard Torque Transfer Wrench Model: TW-MCS01 | ? 1,000 N?m | DAkkS-DKD-R 3-7 Class 0.2 | 228,000/piece | 1. Temperature compensation on zero signal 2. Temperature compensation on span signal 3. Zero balance 4. Measurement and check signals | 10040-20604 | Mechanical Metrology Department | 228000 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor and Connection profile rotatable, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 18,750/range |  | 10040-30310 | Mechanical Metrology Department | 18750 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor and Connection profile rotatable, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-30311 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor rotatable and Connection profile fixed, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-30312 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor rotatable and Connection profile fixed, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 6,550/range |  | 10040-30313 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor fixed and Connection profile rotatable, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 18,750/range |  | 10040-30314 | Mechanical Metrology Department | 18750 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor fixed and Connection profile rotatable, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-30315 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N·m to 1,000 N·m, Sensor and Connection profile fixed, target uncertainty <0.5% | DKD-R 10-8, 0.2% | 12,650/range |  | 10040-30316 | Mechanical Metrology Department | 12650 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Static Torque Calibration devices '- Torque Wrench Calibration Devices | 1 N?m to 1,000 N?m, Sensor and Connection profile fixed, target uncertainty ?0.5% | DKD-R 10-8, 0.2% | 6,550/range |  | 10040-30317 | Mechanical Metrology Department | 6550 | Mechanical Metrology Department | 1 | torque, devices, wrench, calibration, static
| Endurance testing mchaine - Tyre load | At Load index, LI = 0 to 121 (TLCC = 46 kgf to 1,450 kgf) and At Load index, LI = 122 to 279 (TLCC = 1,500 kgf to 136,000 kgf) | 0.047 % for increasing force only 0.35 % for increasing and decreasing force | 12,500/range/direction | UN Regulation No. 30, UN Regulation No. 54, UN Regulation No. 75 and ÁÍ¡. 2718-2560, ÁÍ¡. 2719-2560, ÁÍ¡. 2720-2560 | 10040-30905 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Endurance testing mchaine - Surface speed | 0 km/h to 300 km/h | 0.20 km/h | 6,250 / Direction | UN Regulation No. 30, UN Regulation No. 54, UN Regulation No. 75 and ÁÍ¡. 2718-2560, ÁÍ¡. 2719-2560, ÁÍ¡. 2720-2560 | 10040-30906 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Endurance testing mchaine - Drum diameter | 1,700 mm and 2,000 mm | 0.15 mm | 9350 | UN Regulation No. 30, UN Regulation No. 54, UN Regulation No. 75 and ÁÍ¡. 2718-2560, ÁÍ¡. 2719-2560, ÁÍ¡. 2720-2560 | 10040-30907 | Mechanical Metrology Department | 9350 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Wheel deviation measurement system - Run-out (vertical run-out and horizontal run-out) | 0 mm to ± 2.5 mm | Vertical run-out = 0.50 mm Horizontal run-out = 0.50 mm | 5000 | ÁÍ¡. 2207-2562 | 10040-31001 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rockwell Scale : HRA, HRB, HRC | 10 HRBW to 100 HRBW 20 HRC to 70 HRC | 0.40 HRBW 0.40 HRC |  | by client. | 10050 - 10101 | Mechanical Metrology Department | 7880 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rockwell Scale : HR30TW | 29 HR30TW to 82 HR30TW | 0.40 HR30TW | 9,350/piece |  | 10050 - 10101 | Mechanical Metrology Department | 9350 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vickers : HV according to ISO 6508-3 | HV 5, HV 10, HV 20, HV 30, HV 50, HV 100 |  | 7,250/piece |  | 10050 - 10201 | Mechanical Metrology Department | 7250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Elastomer Hardness Testing Machine | Durometer Type C | ASTM D2240 | 8,550/machine |  | 10050 - 10302 | Mechanical Metrology Department | 8550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Elastomer Hardness Testing Machine | Durometer Type AO | According to ISO48-9, ASTM D2240, JIS K6253, JIS K7215, JIS K6301, DIN 53505 | 8,550/machine |  | 10050 - 10302 | Mechanical Metrology Department | 8550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Elastomer Hardness Testing Machine | Durometer Type D | According to ISO48-9, ASTM D2240, JIS K6253, JIS K7215, JIS K6301, DIN 53505 | 8,550/machine |  | 10050 - 10302 | Mechanical Metrology Department | 8550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Elastomer Hardness Testing Machine | Durometer Type A | According to ISO48-9, ASTM D2240, JIS K6253, JIS K7215, JIS K6301, DIN 53505 | 8,550/machine |  | 10050 - 10302 | Mechanical Metrology Department | 8550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Elastomer Hardness Testing Machine | IRHD-M | According to ISO48-9, ASTM D2240, JIS K6253, JIS K7215, JIS K6301, DIN 53505 | 8,550/machine |  | 10050 - 10302 | Mechanical Metrology Department | 8550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Elastomer Hardness Testing Machine | IRHD Pocket | According to ISO48-9, ASTM D2240, JIS K6253, JIS K7215, JIS K6301, DIN 53505 | 8,550/machine |  | 10050 - 10302 | Mechanical Metrology Department | 8550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Elastomer Hardness Testing Machine | IRHD-N | According to ISO48-9, ASTM D2240, JIS K6253, JIS K7215, JIS K6301, DIN 53505 | 8,550/machine |  | 10050 - 10302 | Mechanical Metrology Department | 8550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Brinell Scale: HB | HBW 5/750 | 1.2% of HBW for Force 7.355 kN (750 kgf) | 6,250/piece |  | 10050 - 10401 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Brinell Scale: HB | HBW 10/500 | 1.2% of HBW for Force 4.903 kN (500 kgf) | 7,800/piece |  | 10050 - 10401 | Mechanical Metrology Department | 7800 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Brinell Scale: HB | HBW 10/3000 | 0.6% of HBW for Force 29.42 kN (3,000 kgf) | 7,800/piece |  | 10050 - 10401 | Mechanical Metrology Department | 7800 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vickers : HV according to ISO 6508-3 | HV 0.1 |  | 7,250/piece |  | 10050 - 10501 | Mechanical Metrology Department | 7250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Certified Reference Block Rockwell Scale : HRA, HRB, HRC | 20 HRA to 95 HRA | 0.40 HRA | 13,670/piece |  | 10050 - 20101 | Mechanical Metrology Department | 13670 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Certified Reference Block Rockwell Scale : HRA, HRB, HRC | 20 HRC to 65 HRC | 0.40 HRC | 13,670/piece |  | 10050 - 20101 | Mechanical Metrology Department | 13670 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Rockwell Scale : HR30TW | 29 HR30TW to 82 HR30TW | 0.40 HR30TW | 17,500/piece |  | 10050 - 20101 | Mechanical Metrology Department | 17500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Certified Reference Block Rockwell Scale : HRA, HRB, HRC | 10 HRBW to 100 HRBW | 0.40 HRBW | 13,670/piece |  | 10050 - 20101 | Mechanical Metrology Department | 13670 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vickers : HV | HV 5, HV 10, HV 20, HV 30, HV 50, HV 100 |  | 25,400/piece |  | 10050 - 20201 | Mechanical Metrology Department | 25400 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CRM | 30 J ± 10 J | According to ISO148-3 | 18,530/set |  | 10050 - 20301 | Mechanical Metrology Department | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CRM | 100 J ± 10 J | According to ISO148-3 | 18,530/set |  | 10050 - 20302 | Mechanical Metrology Department | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Callibration of metellic -charpy pendulum Testing Machine | Low(<40J) and high (?40J)energy | According to ISO148-2 | 81250 |  | 10050 - 20303 | Mechanical Metrology Department | 81250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Brinell Scale: HB | HBW 10/3000 | 0.6% of HBW for Force 29.42 kN (3,000 kgf) | 21,100/piece |  | 10050 - 20401 | Mechanical Metrology Department | 21100 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Brinell Scale: HB | HBW 10/500 | 1.2% of HBW for Force 4.903 kN (500 kgf) | 21,100/piece |  | 10050 - 20401 | Mechanical Metrology Department | 21100 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Brinell Scale: HB | HBW 5/750 | 1.2% of HBW for Force 7.355 kN (750 kgf) | 18,750/piece |  | 10050 - 20401 | Mechanical Metrology Department | 18750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vickers : HV | HV 0.1 |  | 32,800/piece |  | 10050 - 20501 | Mechanical Metrology Department | 32800 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Extensometer | Gauge length up to 200 mm | According to ISO 9513 | 3,910 | /scale | 10050 - 30001 | Mechanical Metrology Department | 3910 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 70 HR15N to 94 HR15N | According to ISO 6508-2 | 6100 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 6100 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 42 HR30N to 86 HR30N | According to ISO 6508-2 | 1250 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 1250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 67 HR15TW to 93 HR15TW | According to ISO 6508-2 | 1250 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 1250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 29 HR30TW to 82 HR30TW | According to ISO 6508-2 | 1250 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 1250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 70 HREW to 100 HREW | According to ISO 6508-2 | 1250 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 1250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 20 HRA to 95 HRA | According to ISO 6508-2 | 10550 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 10550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 10 HRBW to 100 HRBW | According to ISO 6508-2 | 10550 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 10550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Rockwell Hardness Testing Machine - direct and indirect verification - indirect verification - diamond indenter verification - ball indenter verification | 20 HRC to 70 HRC | According to ISO 6508-2 | 4370 | /machine/scale /scale | 10050 - 30102 | Mechanical Metrology Department | 4370 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Scale B | MH01.1 | 12,150 |  | 10050 - 30103 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Tip angle 0°,90°, 136° Line of conjunction 0 to 5 µm | ± 0.15° ± 0.2 µm | 12500 |  | 10050 - 30202 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Calibration of Testing Machine Vickers Hardness Testing Machine | HV 0.2, HV 0.3, HV 0.5, HV 1, HV 5, HV 10, HV 20, HV 30, HV 50 | According to ISO 6507-2 | 10550 |  | 10050 - 30202 | Mechanical Metrology Department | 10550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HV 0.2, HV 0.3, HV 0.5, HV 1, HV 5, HV 10, HV 20, HV 30, HV 50 | According to ISO 6507-2 | 10550 | /machine/scale /scale | 10050 - 30202 | Mechanical Metrology Department | 10550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HV 0.2, HV 0.3, HV 0.5, HV 1, HV 5, HV 10, HV 20, HV 30, HV 50 | According to ISO 6507-2 | 4370 |  | 10050 - 30202 | Mechanical Metrology Department | 4370 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Scale C | MH01.2 | 12,150 |  | 10050 - 30203 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale A, B and C | MH03 | 12,150 |  | 10050 - 30303 | Mechanical Metrology Department | 12150 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HBW 2.5/62.5, HBW 2.5/31.25, HBW 2.5/15.625 | According to ISO 6506-2 | 4370 | /machine/scale /scale | 10050 - 30402 | Mechanical Metrology Department | 4370 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HBW 5/750, HBW 5/125, HBW 5/62.5 | According to ISO 6506-2 | 10550 | /machine/scale /scale | 10050 - 30402 | Mechanical Metrology Department | 10550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of Testing Machine Brinell Hardness testing machine | HBW 10/3000, HBW 10/500, HBW 10/100 | According to ISO 6506-2 | 10550 | /machine/scale /scale | 10050 - 30402 | Mechanical Metrology Department | 10550 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HBW 1/30 | According to ISO 6506-2 | 1250 | /machine/scale /scale | 10050 - 30402 | Mechanical Metrology Department | 1250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HV 5 | MH02.1 | 10,300 |  | 10050 - 30403 | Mechanical Metrology Department | 10300 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Mooney Viscometer | All range | According to ISO 289 | 25,000 | /machine | 10050 - 30602 | Mechanical Metrology Department | 25000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Scale R | MH01.3 | 12,250 | Plastic artifacts shall be provided by client | 10050 - 30803 | Mechanical Metrology Department | 12250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Speed for material testing machines | Speed 0.05 mm/Min to 2,500 mm/Min. | According to | 6,250 | /speed | 10050 - 30902 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Brinell HBW 10/3000 | MH05.1 | 6,850 |  | 10050 - 30903 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Vickers HV 10 | MH04.4 | 6,850 |  | 10050 - 31103 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Vickers HV 0.3 | MH04.1 | 6,850 |  | 10050 - 31203 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HV 10 | MH02.2 | 10,300 |  | 10050 - 31403 | Mechanical Metrology Department | 10300 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | HV 30 | MH02.3 | 10,300 |  | 10050 - 31503 | Mechanical Metrology Department | 10300 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Vickers HV 1 | MH04.2 | 6,850 |  | 10050 - 31603 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Vickers HV 5 | MH04.3 | 6,850 |  | 10050 - 31703 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Scale A | MH01.4 | 34,000 |  | 10050 - 31803 | Mechanical Metrology Department | 34000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Scale 30N | MH01.5 | 34,000 |  | 10050 - 31903 | Mechanical Metrology Department | 34000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Scale 30TW | MH01.6 | 34,000 |  | 10050 - 32003 | Mechanical Metrology Department | 34000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale 15N | MH03.4 | 10,950 |  | 10050 - 32103 | Mechanical Metrology Department | 10950 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale 30N | MH03.5 | 10,950 |  | 10050 - 32203 | Mechanical Metrology Department | 10950 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale A | MH03.1 | 6,850 |  | 10050 - 32303 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale B | MH03.2 | 6,850 |  | 10050 - 32403 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale C | MH03.3 | 6,850 |  | 10050 - 32503 | Mechanical Metrology Department | 6850 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale 15TW | MH03.6 | 10,950 |  | 10050 - 32603 | Mechanical Metrology Department | 10950 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Rockwell Scale 30TW | MH03.7 | 10,950 |  | 10050 - 32703 | Mechanical Metrology Department | 10950 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Vickers HV 0.1 | MH04.5 | 16,700 |  | 10050 - 32803 | Mechanical Metrology Department | 16700 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Vickers HV 0.2 | MH04.6 | 16,700 |  | 10050 - 32903 | Mechanical Metrology Department | 16700 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | Vickers HV 0.5 | MH04.7 | 16,700 |  | 10050 - 33003 | Mechanical Metrology Department | 16700 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2524 | Recalibration of Reference Block Rockwell : HR According to ISO6508-3 : 2023 | HR15N | 0.50 HR15N | 9,350/piece | According to ISO6508-3 : 2023 | 10050-10101 | Mechanical Metrology Department | 9350 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vickers : HV **According to ASTM E92 | HV10, HV 20 |  | 17,500/piece |  | 10050-10206 | Mechanical Metrology Department | 17500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2525 | Testing the nanoindentation hardness of materials | Test force : 0.1 to 480 mN Depth (Dz) : 0 to 7.2 ?m Loading rate : 0.1 mN/min to 10 N/min Testing cycle : 1 to 70 s | 4.0 GPa of Indentation modulus; EIT | 8,500/piece | Based on the ISO 14577-1:2015 | 10050-10601 | Mechanical Metrology Department | 8500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2523 | Certified Reference Block Rockwell : HR According to ISO6508-3 : 2023 | HR15N | 0.50 HR15N | 17,500/piece | According to ISO6508-3 : 2023 | 10050-20101 | Mechanical Metrology Department | 17500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2521 | Vickers : HV According to ISO6507-3 | HV 5, HV 10, HV 20, HV 30, HV 50, HV 100 Range 100 HV to 950 HV | (2.1 + 2600/d2)% of HV | 25,400/piece |  | 10050-20201 | Mechanical Metrology Department | 25400 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vickers : HV According to ASTM E92 | HV10, HV 20 |  | 26,250/piece |  | 10050-20206 | Mechanical Metrology Department | 26250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Calibration of Rockwell standard machine. | Scale A, B, C, E, 30N, 30T, 15N, 15T | According to ISO6508-3 | 32,000 | /machine/scale | 10050-30105 | Mechanical Metrology Department | 32000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Calibration of Rockwell standard machine. | Scale A, B, C, E, 30N, 30T, 15N, 15T |  | 5,620 | /additional scale | 10050-30105 | Mechanical Metrology Department | 5620 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Calibration of Vickers standard machine. | scale HV 5, HV 10, HV 20, HV 30 and HV 50 | According to ISO6507-3 0.08 % for test force 0.1 s for testing cycle 0.4 ?m for measuring system | 32,000 | /machine/scale | 10050-30205 | Mechanical Metrology Department | 32000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Calibration of Vickers standard machine. |  |  | 6,250 | /additional scale | 10050-30205 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Calibration of Brinell standard machine. |  |  | 5,750 | /additional scale | 10050-30305 | Mechanical Metrology Department | 5750 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Calibration of Brinell standard machine. | HBW 10/3000, HBW 10/1500, HBW 10/1000, HBW 10/500, HBW 10/250, HBW 10/100, HBW 5/750, HBW 5/250, HBW 5/125, HBW 5/62.5, HBW 5/25, HBW 2.5/187.5, HBW 2.5/62.5, HBW 2.5/31.25, HBW 2.5/15.625, HBW 2.5/6.25, HBW 1/30, HBW 1/10, HBW1/5, HBW 1/2.5, HBW 1/1 | According to ISO6506-3 | 33,000 | /machine/scale | 10050-30305 | Mechanical Metrology Department | 33000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Calibration of Vickers standard machine. | scale HV 0.2, HV 0.3, HV 0.5 and HV 1 | According to ISO6507-3 0.08 % for test force 0.1 s for testing cycle 0.4 ?m for measuring system | 32,000 | /machine/scale | 10050-30505 | Mechanical Metrology Department | 32000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Calibration of Vickers standard machine. |  |  | 7,100 | /additional scale | 10050-30505 | Mechanical Metrology Department | 7100 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Melt flow index tester (Extrusion plastometer) | All range | According to ASTM D1238 | 15,000 | /machine | 10050-30702 | Mechanical Metrology Department | 15000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Displacement Measuring Systems and Devices Used in Material Testing Machines | Displacement length up to 2,000 mm. | According to ASTM E2309/E2309M-16 | 6,250 | /machine | 10050-30802 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Pressure Balances | Up to 7 MPa | > 5x10-4 x Pe to 1x10-4 x Pe | 50000 | 1. Price per 1 piston & | 10060 - 10101 | Mechanical Metrology Department | 50000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Balances | Up to 7 MPa | < 1x10-4 x Pe | 50000 | cylinder assembly | 10060 - 10102 | Mechanical Metrology Department | 50000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Balances | >7 MPa to 40 MPa | > 5x10-4 x Pe | 55000 | 2. Calibration of weight | 10060 - 10103 | Mechanical Metrology Department | 55000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Balances |  | 5x10-4 x Pe to 1x10-4 x Pe | 55000 | set is included. | 10060 - 10104 | Mechanical Metrology Department | 55000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Balances |  | < 1x10-4 x Pe | 55000 |  | 10060 - 10105 | Mechanical Metrology Department | 55000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Balances | Up to 200 MPa | > 5x10-4 x Pe to 1x10-4 x Pe | 50000 | 1. Price per 1 piston & | 10060 - 10201 | Mechanical Metrology Department | 50000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Balances |  | < 1x10-4 x Pe | 50000 | cylinder assembly | 10060 - 10202 | Mechanical Metrology Department | 50000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Balances | > 200 MPa to 500 MPa | > 5x10-4 x Pe to 2x10-4 x Pe | 60000 | 2. Calibration of weight | 10060 - 10203 | Mechanical Metrology Department | 60000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Balances |  | < 2x10-4 x Pe | 60000 | set is included. | 10060 - 10204 | Mechanical Metrology Department | 60000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | - 1 kPa to 0 kPa | > 6 x 10-3 x Pe | 3,750/range |  | 10060 - 10301 | Mechanical Metrology Department | 3750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | 6x10-3 x Pe to 1x10-3 x Pe | 5,000/range |  | 10060 - 10302 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1x10-3 x Pe, but not smaller than 0,003 mbar | 6,250/range |  | 10060 - 10303 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | 0 kPa to 3 kPa | > 6x10-3 x Pe | 3,750/range |  | 10060 - 10304 | Mechanical Metrology Department | 3750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | 6x10-3 x Pe to 1x10-3 x Pe | 5,000/range |  | 10060 - 10305 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1x10-3 x Pe, but not smaller than 0.003 mbar | 6,300/range |  | 10060 - 10306 | Mechanical Metrology Department | 6300 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | 10 kPa to 7 MPa | > 6x10-3 x Pe | 4,070/range |  | 10060 - 10307 | Mechanical Metrology Department | 4070 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | 1,5 kPa to 7 MPa | 6x10-3 x Pe to 1x10-3 x Pe | 6,250/range |  | 10060 - 10308 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1x10-3 x Pe | 7,500/range |  | 10060 - 10309 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | > 7 MPa to 100 MPa | > 6x10-3 x Pe | 7,040/range |  | 10060 - 10310 | Mechanical Metrology Department | 7040 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | 6x10-3 x Pe to 1x10-3 x Pe | 9,850/range |  | 10060 - 10311 | Mechanical Metrology Department | 9850 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1x10-3 x Pe | 12,670/range |  | 10060 - 10312 | Mechanical Metrology Department | 12670 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | 0 kPa to –100 kPa | Not smaller than 0.2 mbar | 8,750/range |  | 10060 - 10313 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | Not smaller than 1 mbar | 4,070/range |  | 10060 - 10314 | Mechanical Metrology Department | 4070 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | 0 kPa to -15 kPa | > 1x10-3 x Pe | 8,750/range |  | 10060 - 10315 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1x10-3 x Pe | 11,250/range |  | 10060 - 10316 | Mechanical Metrology Department | 11250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | 0 kPa to 15 kPa | > 1x10-3 x Pe | 8,750/range |  | 10060 - 10317 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1x10-3 x Pe | 11,250/range |  | 10060 - 10318 | Mechanical Metrology Department | 11250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) | - 4 kPa to 4 kPa | > 6x10-3 x Pe | 3,750/range |  | 10060 - 10323 | Mechanical Metrology Department | 3750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | 6x10-3 x Pe to 1x10-3 x Pe | 5,000/range |  | 10060 - 10324 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1x10-3 x Pe, but not smaller than 0.003 mbar | 6,300/range |  | 10060 - 10325 | Mechanical Metrology Department | 6300 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (absolute pressure, Pabs) | 80 kPa to 115 kPa | < 1x10-3 x Pabs | 8,750/range |  | 10060 - 10401 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (absolute pressure, Pabs) |  | Not smaller than 0.35 mbar | 4,230/range |  | 10060 - 10402 | Mechanical Metrology Department | 4230 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (absolute pressure, Pabs) | 2 kPa to 7 MPa | Not smaller than 0.35 mbar | 4,230/range |  | 10060 - 10403 | Mechanical Metrology Department | 4230 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (absolute pressure, Pabs) | 1,5 kPa to 40 MPa | > 1x10-4 x Pabs | 8,750/range |  | 10060 - 10404 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (absolute pressure, Pabs) | 0 kPa to 15 kPa | > 1x10-3 x Pabs | 10,000/range |  | 10060 - 10405 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (absolute pressure, Pabs) |  | < 1x10-3 x Pabs | 12,500/range |  | 10060 - 10406 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Mercury Barometers | 85 kPa to 105 kPa | Not smaller than 0.35 mbar | 6,250/range | Cistern type | 10060 - 10501 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Mercury Barometers | 85 kPa to 105 kPa | Not smaller than 0.35 mbar | 7,500/range | Fortin type | 10060 - 10502 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Hydraulic Pressure Measuring Instruments (gauge pressure, Pe) | up to 200 MPa | > 6x10-3 x Pe | 5,000/range |  | 10060 - 10601 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (gauge pressure, Pe) |  | 6x10-3 x Pe to 1x10-3 x Pe | 7,500/range |  | 10060 - 10602 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1 x 10-3 x Pe | 10,000/range |  | 10060 - 10603 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (gauge pressure, Pe) | > 200 MPa to 500 MPa | > 1 x 10-3 x Pe | 8,750/range |  | 10060 - 10604 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (gauge pressure, Pe) |  | < 1 x 10-3 x Pe | 11,250/range |  | 10060 - 10605 | Mechanical Metrology Department | 11250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Differential Pressure at High Static Line ressure (not higher than 40 MPa), DP | 350 kPa maximum differential pressure | 3.5 x 10-5 x DP, but not less than 2.9 Pa | 16,250/sensor head | Price per 1 range of differential pressure and 1 range of high static line pressure | 10060 - 10701 | Mechanical Metrology Department | 16250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Hydraulic Pressure Measuring Instruments (absolute pressure, Pabs) | 0,1 MPa to 200 MPa | > 6x10-3 x Pabs | 6,250/range |  | 10060 - 10801 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (absolute pressure, Pabs) |  | 6x10-3 x Pabs to 1x10-3 x Pabs | 8,750/range |  | 10060 - 10802 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (absolute pressure, Pabs) |  | < 1 x 10-3 x Pabs | 11,250/range |  | 10060 - 10803 | Mechanical Metrology Department | 11250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (absolute pressure, Pabs) | > 200 MPa to 500 MPa | > 1 x 10-3 x Pabs | 10,000/range |  | 10060 - 10804 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Hydraulic Pressure Measuring Instruments (absolute pressure, Pabs) |  | < 1x10-3 x Pabs | 12,500/range |  | 10060 - 10805 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Blood Pressure (Static Mode)(gauge pressure, Pe) | (0 kPa to 40 kPa)(0 mm Hg to 300 mm Hg) | Not smaller than 0.13 kPa ( 1 mm Hg) | 1,875/range | To include leak check | 10060 - 10902 | Mechanical Metrology Department | 1875 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Absolute Gravity |  |  | 120000 | for each additional location | 10060 - 30101 | Mechanical Metrology Department | 120000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Absolute Gravity | 9.75 m/s2 to 9.85 m/s2 | Not smaller than 10 µGal (0.01 ppm) | 240000 | for the first measurement location | 10060 - 30101 | Mechanical Metrology Department | 240000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Local Gravity | 9.75 m/s2 to 9.85 m/s2 | Not smaller than 180 µGal (0.18 ppm) | 40000 | Using relative method | 10060 - 30102 | Mechanical Metrology Department | 40000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Pressure Measuring Instruments (on-site) - (differential gauge pressure,Pe) | 0.1 kPa to 130 kPa | 2.6E-4Pe but not less than 0.03 kPa | 6,250 baht/range | Differential against ambient pressure | 10060 - 30103 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (on-site) - (absolute pressure, Pabs) | 10 kPa to 700 kPa | 2.0E-4Pabs but not less than 0.03 kPa | 6,250 baht/range |  | 10060 - 30103 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Pressure Measuring Device | 0 to 3500 kPa | 0.01%rdg but not less than 0.1 kPa | 6,500/range | Media: Water | 10060-10901 | Mechanical Metrology Department | 6500 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Gas Pressure Measuring Instruments (gauge pressure, Pe) |  | Smaller than 0.2 mbar | 11,250/range |  | 10061 - 10319 | Mechanical Metrology Department | 11250 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Vacuum Gauges (absolute pressure : pirani gauge, thermal conductivity gauge) | 105 Pa to 10-1 Pa | 0.3% rdg to 1% rdg | 7,040/sensor head | Calibration 10 points minimum | 10070 - 10101 | Mechanical Metrology Department | 7040 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vacuum Gauges (absolute pressure : capacitance diaphragm gauge) | 105 Pa to 10-1 Pa | 0.2% rdg to 0.45% rdg | 12,670/sensor head | Calibration 3 decades for the head range (1000, 100,10, 1) torr Calibration 2 decades for the head range 0.1 torr | 10070 - 10201 | Mechanical Metrology Department | 12670 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vacuum Gauges (absolute pressure : hot cathode ion gauge, cold cathode ion gauge) | 10-1 Pa to 10-5 Pa | 0.5% rdg to 6% rdg | 11,090/for first decade +2,250 for next decade |  | 10070 - 10301 | Mechanical Metrology Department | 11090 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vacuum Gauge (absolute pressure : spinning rotor gauge) | 10 Pa to 10-3 Pa | 0.6% rdg to 4.5% rdg | to be determined |  | 10070 - 10401 | Mechanical Metrology Department | 7880 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Additional cost for electrical calibration (which the accuracy is not better than +/-0.01% rdg) |  |  | 2,250/item |  | 10070 - 10501 | Mechanical Metrology Department | 2250 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Adjustment charge of pressure and vacuum devices for permanent laboratory calibration |  |  | 50% of the calibration cost of each range |  | 10070 - 10502 | Mechanical Metrology Department | 50 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Measurement audit report |  |  | 1,000 Baht/report + 1,000 Bath/only for artifact provided + Calibration fee |  | 10070 - 10503 | Mechanical Metrology Department | 1000 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Vacuum Gauges by Static Expansion Method (absolute pressure : capacitance diaphragm gauge, spinning rotor gauge) | 0.13 Pa to 1.3 kPa | 0.50% rdg to 0.25% rdg | 26,400/sensor head |  | 10070 - 10601 | Mechanical Metrology Department | 26400 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Liquid Flowmeter (Turbine Flowmeter) | 0.19 L/min to 150* L/min | 0.075%** | 13,750/ 5 points | Calibration Fluid : Water | 10080 - 10101 | Mechanical Metrology Department | 13750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Liquid Flowmeter (with indicator) | 0.19 L/min to 150* L/min | 0.02 | 7,880/ 5 points | Calibration Fluid : Water | 10080 - 10102 | Mechanical Metrology Department | 7880 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Coriolis Flowmeter (Volumetric method) | 0.2 L/min to 150 L/min | 0.00075 | 13,750/5 points | Calibration Fluid : Water | 10080 - 10103 | Mechanical Metrology Department | 13750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Inter-laboratory comparison on liquid volume flow rate measurement; Artifact : Coriolis flow meter | 25 L/min to 120 L/min | NIMT.MW-L1/2017 | 20,333/range (N = 3) | Calibration Fluid : Water | 10080 - 10105 | Mechanical Metrology Department | 20333 | Mechanical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Inter-laboratory comparison on liquid volume flow rate measurement; Artifact : Coriolis flow meter | 25 L/min to 120 L/min | NIMT.MW-L1/2019 | 15,750/range (N = 2) | Calibration Fluid : Water | 10080 - 10106 | Mechanical Metrology Department | 15750 | Mechanical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Inter-laboratory comparison on water flow meter and pressure of infusion pump analyzer | (20 - 200) mL/h ((0 - 1200) mmHg | 0.35% 0.35 mbar | 12,552.50/range | Infusion device analyzer | 10080 - 10107 | Mechanical Metrology Department | 12552.5 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Liquid Flowmeter (Turbine Flowmeter) | 20 L/min to 1300 L/min | 0.00075 | 13,750 | Calibration Fluid : Water | 10080 - 10108 | Mechanical Metrology Department | 13750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Ultrasonic flow meter (Clamp-on) | 20 L/min to 1300 L/min | 0.01 | 8,750 | Calibration Fluid : Water | 10080 - 10109 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Gas Flowmeter | 0.043 (mg/s) to 500 (mg/s) | 0.0023 | 17,250/range | Calibration Fluid : Nitrogen | 10080 - 10201 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flowmeter with display in standard flow (molbloc-L) | 100 cm3/min to 24000 cm3/min | 0.0025 | 30,000/range | Calibration Gas: Dry Air Reference Conditions @ 0°C, 101.325 kPa | 10080 - 10202 | Mechanical Metrology Department | 30000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| (Mass gas flow rate) | 0.043 (mg/s) to 21545 (mg/s) | 0.0023 | 17,250/range | Calibration Fluid : Dry Air | 10080 - 10202 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flow rate (Gas meter with impluse output) | 1 m3/h to 100 m3/h | 0.002 | 30,000/piece (5 points) | Comparison Method | 10080 - 10203 | Mechanical Metrology Department | 30000 | Mechanical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Gas Flowmeter with display in standard flow (molbloc-S) | 1 m3/h to 65 m3/h | 0.0023 | 30,000/range | Calibration Gas: Dry Air Reference Conditions @ 0°C, 101.325 kPa | 10080 - 10204 | Mechanical Metrology Department | 30000 | Mechanical Metrology Department | 4 | standard, solution, mass, set, ph
| Sonic Nozzle at ambient pressure | 1 m3/h to 100 m3/h | 0.0022 | 25,000/range | Calibration Gas: Ambient Air | 10080 - 10205 | Mechanical Metrology Department | 25000 | Mechanical Metrology Department | 2 | pressure, measuring, instruments, gas, pe
| Orifice flow device (Volumetric method) | (1 to 100) m3/h | 0.01 | 7,500/piece (5-7 points) | Comparison Method | 10080 - 10208 | Mechanical Metrology Department | 7500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Small Water Flow Meter (Mass Mode) | 10 g/h to 50 g/h 50 g/h to 1000 g/h | 0.35% 0.24% | 16,875/range | Calibration Fluid : Water Gravimetric Method | 10080 - 10301 | Mechanical Metrology Department | 16875 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Medium Liquid Flow Meter (Gravimetric Method) | 0.036 kg/min to 0.32 kg/min | 0.003 | 12,500/5 points | Calibration Fluid : Water Gravimetric Method | 10080 - 10302 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Gas Flow rate (Gas meter with flow rate output) | 1.2 kg/h to 117 kg/h | 0.002 | 30,000/piece (5 points) | Comparison Method | 10080 - 10403 | Mechanical Metrology Department | 30000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Liquid Volume (Proving tank: Neck scale type, Weir neck type) | 1 L to 10 L | 0.02% to 0.03% | 6,250/piece | Gravimetric Method | 10080 - 10501 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Liquid Volume (Proving tank: Neck scale type, Weir neck type) | <10 L to 20 L | 0.02% to 0.03% | 10,000/piece | Gravimetric Method | 10080 - 10502 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Liquid Volume (Proving tank: Neck scale type, Weir neck type) | > 2 L to 50 L | 0.0009 | 8,750 / piece | Gravimetric Method | 10080 - 10503 | Mechanical Metrology Department | 8750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Liquid Volume (Proving tank: Neck scale type, Weir neck type) | > 50 L to 200 L | 0.0009 | 13,750/piece | Gravimetric Method | 10080 - 10504 | Mechanical Metrology Department | 13750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Flow speed device | (0.1-40) m/s | 0.005 | 12,500/5 points+2,500/next point | Wind tunnel Calibration Fluid : Air | 10080 - 10701 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Small Water Flow Meter (Volumetric Mode) | 10 mL/h to 50 mL/h 50 mL/h to 1000 mL/h | 0.35% 0.24% | 16,875/range | Calibration Fluid : Water Gravimetric Method | 10080-10104 | Mechanical Metrology Department | 16875 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (1 to 20) L | 0.15% | 6,250 | / piece / range | 10080-10505 | Mechanical Metrology Department | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (>20 to 50) L | 0.15% | 10,000 | / piece / range | 10080-10506 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (>50 to 100) L | 0.15% | 5,000 | / piece / range | 10080-10507 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (>100 to 500) L | 0.15% | 6,875 | / piece / range | 10080-10508 | Mechanical Metrology Department | 6875 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (>500 to 1000) L | 0.15% | 10,000 | / piece / range | 10080-10509 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (>1000 to 2000) L | 0.15% | 13,750 | / piece / range | 10080-10510 | Mechanical Metrology Department | 13750 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (>2000 to 3000) L | 0.15% | 16,250 | / piece / range | 10080-10511 | Mechanical Metrology Department | 16250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of Liquid | (>3000 to 4000) L | 0.15% | 21250 | / piece / range | 10080-10512 | Mechanical Metrology Department | 21250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Ultrasonic Thickness Measurement (Thickness calibration block) | (2.54 - 12.70) mm and 100 mm | 0.3% of reading | 2,500/5 points +500/next point | Direct measurement by ultrasonic thickness system | 10080-10601 | Mechanical Metrology Department | 2500 | Mechanical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
2520 | Flaw Detector | X-axis (25.25 to 170) mm Y-axis (10 to 100) %dB | 0.5% of reading | 10,000/5 points/1 probe+1500/next probe | Direct measurement by ultrasonic system | 10080-10603 | Mechanical Metrology Department | 10000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Laser Doppler Anemometer | 2 m/s to 20 m/s | 0.004 | 30,000/piece | Wind tunnel Calibration Fluid : Air | 10080-10702 | Mechanical Metrology Department | 30000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Wind direction | 0° to 360° | 0.9° | 12,500/5 points +2,500/next point | Wind tunnel Calibration Fluid : Air | 10080-10703 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Air speed anemometers (Cup and Vane anemometer) | (1-30) m/s | 1% of reading | 12,500/15 points +2,500/next point | Wind tunnel Calibration Fluid : Air | 10080-10704 | Mechanical Metrology Department | 12500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Air speed anemometers | (0.1-40) m/s | 0.008 | 10,000/5 points | Wind tunnel Calibration Fluid : Air | 10080-10705 | Mechanical Metrology Department | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2522 | Rain Gauge | (10 - 480) mm/h | 0.01 | 12,500/5 points +2,500/next point | Calibration Fluid : WaterGravimetric Method | 10080-10802 | Mechanical Metrology Department | 12500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flowmeter(Volume gas flow rate) | 2 (cm3/min) to 24000 (cm3/min) | 0.0023 | 17,250/range | 0 degree C, 101.325 kPa | 10081 - 10101 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flowmeter(Volume gas flow rate) | 2 (cm3/min) to 24000 (cm3/min) | 0.0023 | 17,250/range | Reference Conditions @ | 10081 - 10101 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flowmeter(Volume gas flow rate) | 2 (cm3/min) to 24000 (cm3/min) | 0.0023 | 17,250/range | Calibration Fluid : Nitrogen | 10081 - 10101 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flowmeter(Volume gas flow rate) | 2 (cm3/min) to 1 (m3/min) | 0.0023 | 17,250/range | 0 degree C, 101.325 kPa | 10081 - 10102 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flowmeter(Volume gas flow rate) | 2 (cm3/min) to 1 (m3/min) | 0.0023 | 17,250/range | Reference Conditions @ | 10081 - 10102 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Gas Flowmeter(Volume gas flow rate) | 2 (cm3/min) to 1 (m3/min) | 0.0023 | 17,250/range | Calibration Fluid : Dry Air | 10081 - 10102 | Mechanical Metrology Department | 17250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2519 | Diaphragm Gas Meter | (0.06 to 10) m3/h | 0.01 | 5,000 / 3 points (0.06, 2, 10) m3/h / 1 meter +1250/next point | Comparison Method | 10081-10103 | Mechanical Metrology Department | 5000 | Mechanical Metrology Department | 7 | meter, power, phase, ac, wattmeter
76 | Fixed point cells | 0.01 °C | 0.12 mK | 25,000/point | Direct realization | 11010 - 11111 | Thermometry Metrology Department | 25000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
77 | Standard Platinum Resistance Thermometer (SPRT) | -189.3442 oC | 1 mK | 10,000 / point | By ITS-90 Fixed Point Cells, Sheath for temperature above Zn. Minimum sheath length 450 mm | 11010 - 11321 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 4 | standard, solution, mass, set, ph
78 | Semi-Standard Platinum Resistance Thermometer or Platinum Resistance Thermometer | -189.3442 oC | 1 mK | 10,000 / point | By ITS-90 Fixed Point Cells, Sheath for temperature above Zn. Minimum sheath length 450 mm | 11010 - 11322 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
79 | Standard Platinum Resistance Thermometer (SPRT) | -189.3442 oC to 0.01 oC | 1.0 mK | 10,000 / point | By ITS-90 Fixed Point Cells, Sheath for temperature above Zn. Minimum sheath length 450 mm | 11010 - 11323 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 4 | standard, solution, mass, set, ph
80 | Semi-Standard Platinum Resistance Thermometer or Platinum Resistance Thermometer | -189.3442 oC to 0.01 oC | 1.0 mK | 10,000 / point | By ITS-90 Fixed Point Cells, Sheath for temperature above Zn. Minimum sheath length 450 mm | 11010 - 11324 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
2517 | Fixed point cells | 419.527 oC | 0.6 mK | 25,000/point | Direct realization | 11010-1111 | Thermometry Metrology Department | 25000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
124 | Calibration of Gallium cell: model A | 29.7646 °C | 5 mK | 12,500 /point | Direct measurement with SPRT | 11020 - 12111 | Thermometry Metrology Department | 12500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
125 | Calibration of Gallium cell: model B | 29.76 °C | 20 mK | 6,500 /point | Direct measurement with PRT | 11020 - 12112 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
159 | Standard Platinum Resistance Thermometer (SPRTs) | -196oC | 25 mK | 5,000 first point + 1,500/next point At least 6 points | By Comparison with SPRTs in stirred liquid bath Minimum sheath length 300 mm | 11020 - 12221 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 4 | standard, solution, mass, set, ph
109 | Standard Platinum Resistance Thermometer (SPRTs) | -196oC | 16 mK | 6,500 first point + 1,500/next point At least 6 points | By Comparison with SPRTs in stirred liquid bath Minimum sheath length 300 mm | 11020 - 12221 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 4 | standard, solution, mass, set, ph
110 | Industrial Platinum Resistance Thermometers (IPRTs) | -196oC | 16 mK | 6,500 first point + 1,500/next point At least 6 points | By Comparison with SPRTs in stirred liquid bath Minimum sheath length 300 mm | 11020 - 12222 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
160 | Industrial Platinum Resistance Thermometers (IPRTs) | -196oC | 25 mK | 5,000 first point + 1,500/next point At least 6 points | By Comparison with SPRTs in stirred liquid bath Minimum sheath length 300 mm | 11020 - 12222 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
112 | Noble-metal Thermocouples (Type S, Type R, Type B) | 231.928 °C | 0.1 K | 7500 bath/point | By ITS-90 Fixed Point Cells | 11020 - 12311 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
114 | Noble-metal Thermocouples (Type S, Type R, Type B, Au/Pt) | -40 °C to 0 °C | 0.2 K | 4,000 first point + 1,200/next point | By comparison with standard TC or SPRT in comparison furnace Minimum sheath length 450 mm | 11020 - 12312 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
115 | Base-metal Thermocouples | -196 °C | 0.3 K | 4,000 first point | By comparison with standard TC | 11020 - 12321 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
113 | Pure-metal thermocouples | 231.928 °C | 0.05 K | 8,000 bath/point for Sn, Zn, Al and 10,000 bath for Ag, Cu, Co-C | By ITS-90 Fixed Point Cells / Co-C fixed point cell (At least 3 points) | 11020 - 12331 | Thermometry Metrology Department | 8000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
117 | Liquid-in-glass Thermometer | -40 °C to 90 °C | 15 mK | 4,000 first point+ 1,000/next point | By Comparison with SPRTs in stirred liquid bath | 11020 - 12411 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
118 | Digital Thermometer with resistance probe | - 196 oC | 15 mK | 5,000 first point | By Comparison with SPRTs in stirred liquid bath/ Furnace/ LN2 comparator | 11020 - 12711 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 5 | digital, dna, analog, plasmid, dial
161 | Digital Thermometer (with resistance probe) | - 196 oC | 40 mK | 4,000 first point + 1,500 / next point | By Comparison with SPRTs in stirred liquid bath/ Furnace/ LN2 comparator | 11020 - 12712 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 5 | digital, dna, analog, plasmid, dial
122 | Surface Thermomter | 35 º C to 200 º C | 0.5 K | 5,000 first point + 1,500/next point | Comparison with its reading with the temperature of reference surface heat source. | 11020 - 12713 | Thermometry Metrology Department | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
162 | Digital Thermometer with thermocouple probe | - 196 °C | 0.6 K | 4,000 first point+ 1,500/next point | By Comparison with SPRTs /TC in stirred liquid bath /Furnace | 11020 - 12714 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 5 | digital, dna, analog, plasmid, dial
119 | Digital Thermometer with thermocouple probe | - 196 oC | 0.4 K | 5,000 first point+ 1,500/next point | By Comparison with SPRTs /TC in stirred liquid bath /Furnace | 11020 - 12714 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 5 | digital, dna, analog, plasmid, dial
120 | Temperature Bath | -80oC to 25oC | 20 mK | 4,000 first point + 1,000/next point | By Comparison with SPRTs | 11020 - 12861 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
121 | Temperature block calibrator and Furnace | -40 °C to 660 °C | 0.1 K | 5,000 first point + 3,500/next point | By Comparison with two SPRTs/TCs in measurement zone of temperature block calibrator | 11020 - 12862 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
123 | Surface calibrator | 35 º C to 200 º C | 0.4 K | 5,000 first point + 2,000/next point | Comparison with reference standard temperature | 11020 - 12863 | Thermometry Metrology Department | 5000 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
116 | Fabrication of Standard Thermocouples Type R, Type S, Type B with fixed point calibration (wires not included) | 0 °C to 1100 °C | 0.3 K | 36,500 | Electrical Annealing and re-fabrication thermocouple with calibration certificate by ITS-90 Fixed Point Cells. | 11020 - 22314 | Thermometry Metrology Department | 36500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
111 | Thermistors | -80 º C to 100 º C | 10 mK | 3,000 first point + 1,500/next point | Comparison with SPRTs in stirred liquid baths | 11020-12231 | Thermometry Metrology Department | 3000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2515 | Temperature transmitter | -50 oC to 200 oC | 1.5 K | 5,000 first point + 1,500/next point | y Comparison with SPRTs in stirred liquid bath | 11020-12621 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
152 | Supply of Gallium Cell model A with Calibration | 29.7646 °C | 5 mK | 32,000 | Ga cell suitable for IPRT, DTM calibration & verification | 11020-22111 | Thermometry Metrology Department | 32000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
153 | Supply of Gallium Cell model B with Calibration | 29.76 °C | 20 mK | 12,000 | Ga cell suitable for DTM, IOT sensors calibration & verification | 11020-22112 | Thermometry Metrology Department | 12000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
149 | Supply of Thermocouple Type S with fixed point calibration | 0 °C to 1100 °C | 0.3 K | 99,000 | Supply of thermocouple with calibration certificate by ITS-90 Fixed Point Cells. | 11020-22311 | Thermometry Metrology Department | 99000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
150 | Supply of Thermocouple Type S with fixed point calibration | 0 °C to 1100 °C | 0.3 K | 102,500 | Supply of thermocouple with calibration certificate by ITS-90 Fixed Point Cells. | 11020-22311 | Thermometry Metrology Department | 102500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
151 | Supply of Thermocouple Type B with fixed point calibration | 0 °C to 1100 °C | 0.3 K | 102,500 | Supply of thermocouple with calibration certificate by ITS-90 Fixed Point Cells. | 11020-22313 | Thermometry Metrology Department | 102500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
81 | Fixed-point Blackbody (Primary and Secondary Fixed-point Black body) | 1084.62 oC (ITS-90) | 0.14 K | 24,000 / point | Comparison to Standard Fixed-point Blackbody using 0.65 µm standard radiation thermometer (ITS-90 or thermodynamic temperature) | 11030 - 11121 | Thermometry Metrology Department | 24000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
82 | Fixed-point Blackbody (Primary and Secondary Fixed-point Black body) | 1 084.62 oC (ITS-90) | 0.14 K | 24,000 / point | Comparison to Standard Fixed-point Blackbody using 0.9 µm standard radiation thermometer (ITS-90 or thermodynamic temperature) | 11030 - 11122 | Thermometry Metrology Department | 24000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
83 | Fixed-point Blackbody (Primary and Secondary Fixed-point Black body) | 1 084.62 oC | 0.19 K | 24,000 / point | Comparison to Standard Fixed-point Blackbody using 1.6 µm standard radiation thermometer | 11030 - 11123 | Thermometry Metrology Department | 24000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
84 | Optical Pyrometer/Standard radiation thermometer | 156.5985 °C to 1084.60 °C | 0.14 K to 0.40 K | 12,000/point | By Comparison with fixed pont blackbodies (ITS-90) | 11030 - 11411 | Thermometry Metrology Department | 12000 | Thermometry Metrology Department | 4 | standard, solution, mass, set, ph
85 | 0.9 µm radiation thermometer | 960 °C to 2000 °C | 0.2 K to 1.1 K | 48,000/point | Direct realisation by measurement of | 11030 - 11412 | Thermometry Metrology Department | 48000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
86 | 0.65 µm radiation thermometer | 960 °C to 2500 °C | 0.2 K to 1.4 K | 48,000/point | Direct realisation by measurement of | 11030 - 11413 | Thermometry Metrology Department | 48000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
98 | Variable temperature blackbody sources for infrared ear/forehead thermometers | 22 º C to 42 ºC | 44 mK | 7,500 first point + 2,500/next point | Comparison with standard liquid bath blackbodies through infrared ear thermometers | 11030 - 12524 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
100 | Radiation Thermometer/Infrared Thermometer/Thermal imaging Camera | -40 °C to 15 °C | 0.15 K to 0.30 K | 5,000 /point* | By Comparison with PRT or Transfer Standard Radiation Thermometer | 11030 - 12541 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
101 | Radiation Thermometer/Infrared Thermometer/Thermal imaging Camera | >15 °C to < 1100 °C | 0.15 K to 3.0 K | 7,500 first point+2,000/next point | By Comparison with Transfer Standard Radiation Thermometer or PRT | 11030 - 12542 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
103 | Radiation Thermometer/Infrared Thermometer/Thermal imaging Camera | 1100 °C to 2500 °C | 2 K to 5 K | 8,500 first point+ 3,000/next point | By Comparison with Transfer Standard Radiation Thermometer | 11030 - 12543 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
104 | Infrared ear thermometers (Resolution ³ 0.1 K) | 35 ºC to 42 ºC | 0.2 K | 3,750 first point + 1,000/next point | Comparison with PRT in liquid bath blackbody | 11030 - 12544 | Thermometry Metrology Department | 3750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
105 | Infrared ear thermometer (Resolution < 0.1 K) | 35 ºC to 42 ºC | 0.10 K | 3,750 first point + 1,000/next point | Comparison with PRT in liquid bath blackbody | 11030 - 12545 | Thermometry Metrology Department | 3750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
102 | Radiation Thermometer/Infrared Thermometer/Thermal imaging Camera | 300 °C to 600 °C | 0.45 K to 0.80 K | 7,500 first point+4,500/next point | By Comparison with PRT or Transfer Standard Radiation Thermometer | 11030 - 12546 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
99 | Effective emissivity for variable temperature blackbody cavity (temperature range 22 °C to 42 °C) | 0 to 1 | 0.037 to 0.0025 | 2,500 | Comparison with standard blackbody cavity by measurement of temperature difference | 11030 - 16311 | Thermometry Metrology Department | 2500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
87 | 650 nm / 900 nm Standard radiation thermometer | 1084.6721 °C | 0.16 K to 0.22 K | 12,000 | By Comparison with Cu fixed point blackbody (Thermodynamic temperature) | 11030 - 17121 | Thermometry Metrology Department | 12000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
88 | 650 nm radiation thermometer | 960 °C to 2500 °C | 0.20 K to 1.45 K | 48,000 | Direct realisation by measurement of spectral responsivity (Thermodynamic temperature) | 11030 - 17122 | Thermometry Metrology Department | 48000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
93 | Variable Temperature Blackbody Radiation Sources | 300 ºC to 600 ºC | 0.46 K to 0.80 K | 8,500 first point + 6,250/next point | Comparison with reference blackbodies through infrared thermometer spectral wavelength 8 mm to 14 mm | 11030-12520 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
89 | Variable Temperature Blackbody Radiation Sources | 0 ºC to < 15 ºC | 0.4 K | 7,500 /point | Comparison with reference blackbodies through infrared thermometer spectral wavelength 8 mm to 13 mm | 11030-12521 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
94 | Variable Temperature Blackbody Radiation Sources | 150 ºC to 1 100 ºC | 0.7 K to 1.6 K | 8,500 first point + 3,750/next point | Comparison with 1.6 mm standard radiation thermometer | 11030-12522 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
95 | Variable Temperature Blackbody Radiation Sources | 1 000 ºC to 2 500 ºC | 1.6 K to 2.8 K | 8,500 first point + 3,750/next point | Comparison with 650 nm standard radiation thermometer | 11030-12523 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
90 | Variable Temperature Blackbody Radiation Sources | 15 ºC to 500 ºC | 0.4 K to 2.1 K | 8,500 first point + 3,750/next point |  | 11030-12525 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
91 | Variable Temperature Blackbody Radiation Sources | -40 ºC to < 15 ºC | 0.5 K to 0.2 K | 7,500 /point | Comparison with reference blackbodies through infrared thermometer spectral wavelength 8 mm to 14 mm | 11030-12526 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
92 | Variable Temperature Blackbody Radiation Sources | 15 ºC to 300 ºC | 0.2 K to 0.46 K | 8,500 first point + 3,750/next point | Comparison with reference blackbodies through infrared thermometer spectral wavelength 8 mm to 14 mm | 11030-12527 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
96 | Variable Temperature Blackbody Radiation Sources | 420 ºC to 1 600 ºC | 0.8 K to 2.0 K | 8,500 first point + 3,750/next point | Comparison with 900 nm standard radiation thermometer | 11030-12528 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
97 | Variable Temperature Blackbody Radiation Sources | 400 ºC to 1 000 ºC | 1.2 K to 2.0 K | 8,500 first point + 3,750/next point | Comparison with 8 mm to 14 mm standard radiation thermometer | 11030-12529 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
2516 | Fixed-point Blackbody (thermodynamic temperature) | Zn (419.5408 oC) | 0.14 K | 24,000 / point | Comparison to Standard Fixed-point Blackbody using 0.9 ?m standard radiation thermometer (thermodynamic temperature) | 11030-17112 | Thermometry Metrology Department | 24000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
2518 | Metal-Carbon Fixed point cells (thermodynamic temperature) | Re-C (2474 oC) | 1.6 K | 60,000/point | Comparison with Ag/Cu fixed point using a 650-nm radiation thermometers (thermodynamic temperature) | 11030-17113 | Thermometry Metrology Department | 60000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
2514 | 650 nm standard radiation thermometer (thermodynamic temperature) | 2474 ?C | 1.7 K | 48,000/point | By Comparison with Metal-Carbon fixed point blackbody cells (Thermodynamic temperature) | 11030-17123 | Thermometry Metrology Department | 48000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
145 | Blackbody cavity (En-type) | 35.0 °C to 42.0 °C | Emissivity : 0.9995 ± 0.0005 | 26000 | Model: NIMT-TM-BC-01En | 11030-32521 | Thermometry Metrology Department | 26000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
146 | Blackbody cavity | 30.0 °C to 40.0 °C | Emissivity : 0.999 ± 0.001 | 45000 | Model: NIMT-TM-BC-02FHT (aperture 50 mm suitable for infrared/forehead thermometers) | 11030-32522 | Thermometry Metrology Department | 45000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
147 | Reference thermometer (1 probe) | -50 °C to 70 °C | 20 mK | 10300 | Model: NIMT-TM-DTM-01 | 11030-32711 | Thermometry Metrology Department | 10300 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
148 | Reference thermometer (2 probes) | -50 °C to 70 °C | 20 mK | 12700 | Model: NIMT-TM-DTM-02 | 11030-32712 | Thermometry Metrology Department | 12700 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
142 | Reference temperature source for testing / calibration of clinical thermometers | 35.0 °C to 42.0 °C (up to 150 °C) | 40 mK (uniformity and stability) | 81400 | Model: NIMT-TM-B-01 | 11030-32861 | Thermometry Metrology Department | 81400 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
143 | Reference temperature source for testing / calibration of clinical thermometers | 35.0 °C to 42.0 °C (up to 150 °C) | 40 mK (uniformity and stability) | 55700 | Model: NIMT-TM-B-02 | 11030-32862 | Thermometry Metrology Department | 55700 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
144 | Reference temperature source for testing / calibration of clinical thermometers | 30.0 °C to 40.0 °C (up to 150 °C) | 30 mK (uniformity and stability) | 105000 | Model: NIMT-TM-B43L-03 (volume 43 Liters) | 11030-32863 | Thermometry Metrology Department | 105000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
138 | Air temperature sensor and Relative Humidity Sensor | -70 oC to 60 oC | 0.1 oC | 5,000/First point +2,500 Next point | Comparison to a reference thermometer in an air circulated chamber immersed in liquid bath | 11040 - 12611 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
127 | Dew Point Hygrometer | - 75 °C to -25 °C | 0.3 K | 8,000/point | Comparison with two | 11040 - 13111 | Thermometry Metrology Department | 8000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
128 | Dew Point Hygrometer | - 25 °C to -10 °C | 0.2 K | 8,000/point | Comparison with two | 11040 - 13112 | Thermometry Metrology Department | 8000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
129 | Dew Point Hygrometer | -10 °C to 1 °C | 0.07 K | 7,000/point | Comparison with two pressure | 11040 - 13113 | Thermometry Metrology Department | 7000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
130 | Dew Point Hygrometer | 1 °C to 30 °C | 0.06 K | 7,000/point | Comparison with two pressure | 11040 - 13114 | Thermometry Metrology Department | 7000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
131 | Dew Point Hygrometer | 30 oC to 60 oC | 0.07 K | 8,000/point | Comparison with two pressure humidity generator | 11040 - 13115 | Thermometry Metrology Department | 8000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
126 | Dew Point Hygrometer | -90 oC to -75 oC | 0.4 K | 8,000/point | Comparison with two | 11040 - 13116 | Thermometry Metrology Department | 8000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
132 | Dew Point Hygrometer | 10 °C to 95 °C | 0.03 K | 10,000/point | Comparison with NIMT single temperature humidity generator | 11040 - 13116 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
133 | Relative Hygrometer(At chamber temperature : 15 oC to 60 oC ) | 15 % to 50 % | 0.40% | 8,000/point | Comparison with two temperature two pressure Humidity Generator | 11040 - 13311 | Thermometry Metrology Department | 8000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
166 | Relative Hygrometer(At chamber temperature : 15 oC to 60 oC ) | 15 % to 50 % | 0.60% | 6,300/point | Comparison with two temperature two pressure Humidity Generator | 11040 - 13311 | Thermometry Metrology Department | 6300 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
167 | Relative Hygrometer(At chamber temperature : 15 oC to 60 oC ) | 50 % to 80 % | 1.00% | 5,630/point | Comparison with two temperature two pressure Humidity Generator | 11040 - 13312 | Thermometry Metrology Department | 5630 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
134 | Relative Hygrometer(At chamber temperature : 15 oC to 60 oC ) | 50 % to 80 % | 0.70% | 7,000/point | Comparison with two temperature two pressure Humidity Generator | 11040 - 13312 | Thermometry Metrology Department | 7000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
135 | Relative Hygrometer(At chamber temperature : 15 oC to 60 oC ) | 80 % to 98 % | 0.90% | 7,000/point | Comparison with two temperature two pressure Humidity Generator | 11040 - 13313 | Thermometry Metrology Department | 7000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
168 | Relative Hygrometer(At chamber temperature : 15 oC to 60 oC ) | 80 % to 98 % | 1.20% | 5,630/point | Comparison with two temperature two pressure Humidity Generator | 11040 - 13313 | Thermometry Metrology Department | 5630 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
136 | Thermo-Hygrometer | Humidity (35 to 95)%RH | (1.5 to 3) %RH | 3,130 first point + 500/next point | Comparison with PRT and Chilled Mirror Hygrometer in Controlled Chamber | 11040 - 13411 | Thermometry Metrology Department | 3130 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
137 | Data Logger (Air Type) | Temperature 15 °C to 50 °C | (0.5 to 1) °C | 3,130 first point+ 500/next point | Comparison with PRT and Chilled Mirror Hygrometer in Controlled Chamber | 11040 - 13412 | Thermometry Metrology Department | 3130 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
139 | Dew-Point Generator | (-10 oC to 50 oC) dew-point | 0.07 oC to 0.1 oC | 17,500/First point +4,000 Next point | Comparison with dew-point hygrometer | 11040 - 14411 | Thermometry Metrology Department | 17500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
154 | CRM for Moisture Content in Rice or Paddy | 10 %MC to 20 %MC | 0.15 %MC to 0.2 %MC | 2,500 Bath/100 g/ %MC | The assigned value was obtained from loss on drying method. | 11040 - 25111 | Thermometry Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
155 | CRM for Moisture Content in Rice Powder | 5% %MC to 20 %MC | 0.15 %MC to 0.3 %MC | 2,500 Bath/100 g/ %MC | The assigned value was obtained from loss on drying method. | 11040 - 25112 | Thermometry Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
156 | CRM for Saturated Salt Solution for relative Humidity Unit | Humidity (33 to 75)%RH at Air Temperature (15 °C to 25 °C) | (2.0 to 3.0) %RH | 2,500 Bath/1 Bottle/ %RH | The assigned value was obtained from a thermo-hygrometer in the relative humidity (RH) fixed points method. | 11040 - 25113 | Thermometry Metrology Department | 2500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
157 | Temperature Measurement Platform | -30 oC to 100 oC | 30 mK | 2,500 Bath/Hr. | Comparison with temperature sensor standard in liquid calibration bath. | 11040 - 25114 | Thermometry Metrology Department | 2500 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
158 | Humidity Measurement Platform | 33%RH to 75%RH | (1.5 to 2.5)%RH | 2,500 Bath/Hr. | Comparison with Humidity sensor Standard and humidity ambient by Saturated salt solutions | 11040 - 25115 | Thermometry Metrology Department | 2500 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
140 | Relative humidity generator | 12 % to 92 % (at the air temperature from 10 °C to 40 °C) | 0.1 % to 0.6 % | 17,500/First point +4,000 Next point | Comparison to dew-point hygrometer and reference thermometer | 11040-14421 | Thermometry Metrology Department | 17500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
141 | Moisture meter in rice and paddy (Capacitance, Resistance and NIR type) | 10 %MC to 30 %MC | 0.5 %MC | 1,500/point | Comparison with CRM moisture Content in paddy | 11040-15111 | Thermometry Metrology Department | 1500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
163 | Dew Point Hygrometer | -10 °C to 1 °C | 0.1 K | 5,630/point | Comparison with two pressure | 11041 - 13113 | Thermometry Metrology Department | 5630 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
164 | Dew Point Hygrometer | 1 °C to 30 °C | 0.1 K | 5,630/point | Comparison with two pressure | 11041 - 13114 | Thermometry Metrology Department | 5630 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
165 | Dew Point Hygrometer | 30 oC to 60 oC | 0.1 K | 6,300/point | Comparison with two pressure humidity generator | 11041 - 13115 | Thermometry Metrology Department | 6300 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
108 | Thermal conductivity of solid insulation materials | (0 to 2) W/m·K temperature 25 °C | 2% | 35,000 /point | Measurement by Guarded Hot Plate (GHP) Method | 11050-16111 | Thermometry Metrology Department | 35000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
106 | Spectral emittance (emissivity) (non-IR transparent material) | 0 < E < 1 1 500 nm to 2 500 nm | 6 % to 9.5 % | 3,750 /sample | Comparison with gold-coated reference standard | 11050-16311 | Thermometry Metrology Department | 3750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
107 | Spectral emittance (emissivity) (non-IR transparent material) | 0.9 ? E < 1 5 mm to 20 mm | 6% | 5,000 /sample | Comparison with gold-coated reference standard | 11050-16312 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
132 | Sucrose Standard for Optical Measurements | Purity 997.4 mg/g Cell length 100 mm Polarization at 546 nm: 99.96 (°Z) Optical rotation at 546 nm: 20.38 (degrees) Specific rotation at 546 nm: 78.34 Polarization at 589 nm: 99.95 (°Z) Optical rotation at 589 nm:17.31 (degrees) Specific rotation at 589 nm: 66.52 Cell length 200 mm Polarization at 546 nm: 99.94 (°Z) Optical rotation at 546: 40.75 (degrees) Specific rotation at 546: 99.94 Polarization at 589 nm: 99.89 (°Z) Optical rotation at 589 nm: 34.59 (degrees) Specific rotation at 589 nm: 66.47 | 1.1 mg/g 0.62 (ºZ) 0.13 (degrees) 0.49 0.62 (ºZ) 0.13 (degrees) 0.49 0.60 (ºZ) 0.11 (degrees) 0.43 0.62 (ºZ) 0.25 (degrees) 0.49 0.61 (ºZ) 0.21 (degrees) 0.42 | 6,000/bag |  | 12011-20101 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
103 | 10% Brix Sucrose Standard Solution | 10.00 % Brix, 1.34782 nD | 0.01 % Brix, 0.00007 nD | 3,600/15 mL |  | 12011-20202 | Chemical Metrology Department | 3600 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
104 | 20% Brix Sucrose Standard Solution | 20.00 % Brix, 1.36384 nD | 0.01 % Brix, 0.00007 nD | 3,600/15 mL |  | 12011-20203 | Chemical Metrology Department | 3600 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
105 | 30% Brix Sucrose Standard Solution | 30.00 % Brix, 1.38115 nD | 0.01 % Brix, 0.00007 nD | 3,600/15 mL |  | 12011-20204 | Chemical Metrology Department | 3600 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
106 | 50% Brix Sucrose Standard Solution | 50.00 % Brix, 1.42006 nD | 0.01 % Brix, 0.00008 nD | 3,600/15 mL |  | 12011-20205 | Chemical Metrology Department | 3600 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
107 | 60% Brix Sucrose Standard Solution | 60.00 % Brix, 1.44190 nD | 0.01 % Brix, 0.00007 nD | 3,600/15 mL |  | 12011-20206 | Chemical Metrology Department | 3600 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
108 | Benzene in methanol | 1,000 mg/L | 1.4 % relative | 4,750/8 mL |  | 12011-20207 | Chemical Metrology Department | 4750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
109 | Ethylbenzene in methanol | 1,000 mg/L | 1.1 % relative | 4,750/8 mL |  | 12011-20208 | Chemical Metrology Department | 4750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
110 | Toluene in methanol | 1,000 mg/L | 1.4 % relative | 4,750/8 mL |  | 12011-20209 | Chemical Metrology Department | 4750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
111 | o-xylene in methanol | 1,000 mg/L | 2.2 % relative | 4,750/8 mL |  | 12011-20210 | Chemical Metrology Department | 4750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
112 | m-xylene in methanol | 1,000 mg/L | 1.0 % relative | 4,750/8 mL |  | 12011-20211 | Chemical Metrology Department | 4750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
113 | p-xylene in methanol | 1,000 mg/L | 1.3% relative | 4,750/8 mL |  | 12011-20212 | Chemical Metrology Department | 4750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
114 | Organochlorine Pesticide Mix set 1 | 100 mg/kg | 3% relative | 11,600/1.1 mL |  | 12011-20213 | Chemical Metrology Department | 11600 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
115 | Organophosphate Pesticide Mix set 1 | 100 mg/kg | 5% relative | 7,700/1.1 mL |  | 12011-20214 | Chemical Metrology Department | 7700 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
116 | Organophosphate Pesticide Mix set 2 | 100 mg/kg | 3% relative | 13,300/1.1 mL |  | 12011-20215 | Chemical Metrology Department | 13300 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
117 | Acaricide Pesticide Mix set 1 | 100 mg/kg | 3% relative | 7,000/1.1 mL |  | 12011-20216 | Chemical Metrology Department | 7000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
118 | Carbamate Pesticide Mix set 1 | 100 mg/kg | 3% relative | 8,100/1.1 mL |  | 12011-20217 | Chemical Metrology Department | 8100 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
119 | Pyrethroid Pesticide Mix set 1 | 100 mg/kg | 2 % relative | 11,000/1.1 mL |  | 12011-20218 | Chemical Metrology Department | 11000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
120 | Organophosphate Pesticide Mix set 3 | 100 mg/kg | 4% relative | 14,500/1.1 mL |  | 12011-20219 | Chemical Metrology Department | 14500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
121 | Organophosphate Pesticide Mix set 4 | 100 mg/kg | 4% relative | 11,000/1.1 mL |  | 12011-20220 | Chemical Metrology Department | 11000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
122 | Organochlorine Pesticide Mix set 2 | 100 mg/kg | 2% relative | 16,000/1.1 mL |  | 12011-20221 | Chemical Metrology Department | 16000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
123 | Pyrethroid Pesticide Mix set 2 | 100 mg/kg | 3 % relative | 10,000/1.1 mL |  | 12011-20222 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
124 | Organophosphate Pesticide Mix set 5 | 100 mg/kg | 3% relative | 12,100/1.1 mL |  | 12011-20223 | Chemical Metrology Department | 12100 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
125 | Acaricide Pesticide Mix set 2 | 100 mg/kg | 2 % relative | 14,000/1.1 mL |  | 12011-20224 | Chemical Metrology Department | 14000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
126 | 0.01% Brix Sucrose Standard Solution | 0.01% | 2.0% relative | 3,700/ 15mL |  | 12011-20226 | Chemical Metrology Department | 3700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
127 | 0.02% Brix Sucrose Standard Solution | 0.02% | 1.4% relative | 3,700/ 15mL |  | 12011-20227 | Chemical Metrology Department | 3700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
128 | 0.03% Brix Sucrose Standard Solution | 0.03% | 0.5% relative | 3,700/ 15mL |  | 12011-20228 | Chemical Metrology Department | 3700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
129 | 0.04% Brix Sucrose Standard Solution | 0.04% | 0.6% relative | 3,700/ 15mL |  | 12011-20229 | Chemical Metrology Department | 3700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
130 | 0.05% Brix Sucrose Standard Solution | 0.05% | 0.8% relative | 3,700/ 15mL |  | 12011-20230 | Chemical Metrology Department | 3700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
131 | 40% Brix Sucrose Standard Solution | 40.00% Brix, 1.39986 nD | 0.01 %Brix, 0.00007 nD | 3,600/ 15mL |  | 12011-20231 | Chemical Metrology Department | 3600 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
133 | Ethanol-Water Standard Solution (0.45 mg/g) | 0.45 mg/g | 1% Relative | 1,000/bottle |  | 12011-20233 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
134 | Ethanol-Water Standard Solution (0.55 mg/g) | 0.55 mg/g | 1% Relative | 1,000/bottle |  | 12011-20234 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
135 | Ethanol-Water Standard Solution (1.0 mg/g) | 1.00 mg/g | 1% Relative | 1,000/bottle |  | 12011-20235 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
136 | Aflatoxin B1 Standard Solution | 2.74 µg/g | 4% relative | 3,000/bottle |  | 12011-20236 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
137 | Aflatoxin B2 Standard Solution | 2.72 µg/g | 8% relative | 3,000/bottle |  | 12011-20237 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
138 | Aflatoxin G1 Standard Solution | 2.87 ?g/g | 6 % relative | 3,000/bottle |  | 12011-20238 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
139 | Aflatoxin G2 Standard Solution | 2.69 ?g/g | 7 % relative | 3,000/bottle |  | 12011-20239 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
140 | Sodium Chloride Standard Solution | 0.30 g/100g, 5.84 g/100 g, 12.50 g/100 g | 2% relative | 1,500/bottle |  | 12011-20240 | Chemical Metrology Department | 1500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
141 | High purity pinostrobin standard | 999.5 mg/g | 1% relative | 10,000/bottle | 50 mg | 12011-20241 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
142 | Pesticide Standard Solution (Custom-made) | 100mg/kg | 2.5% relative | 10,000-60,000/5 ampoules x1mL |  | 12011-20242 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
143 | Sucrose Standard Solution (Custom-made) | 5-60% brix | 0.01% brix | 3,700/15 mL |  | 12011-20243 | Chemical Metrology Department | 3700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
144 | Ethanol standard | 939.9 mg/g | 0.6 % relative | 1,000/ampoule | 15 mL | 12011-20244 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
145 | Impurities in ethanol | water content: 59.81 mg/g | 1% relative | 800/bottle | 150 mL | 12011-20245 | Chemical Metrology Department | 800 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
146 | BTEX mixture in Methanol | Benzene: 977 ug/g | 6% relative | 8,000/ ampoule |  | 12011-20246 | Chemical Metrology Department | 8000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
147 | High purity andrographolide standard | 994.0 mg/g | 0.5% relative | 5,000/bottle | 50 mg | 12011-20247 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
148 | High purity panduratin A standard | 991.8 mg/g | 0.6% relative | 20,000/bottle | 10 mg | 12011-20248 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2570 | Ethanol standard solution (95%ABV) (TRM-S-5049a) | Density (in air) at 20 oC (kg/m3) = 809.01, Alcoholic Strength (% ABV) = 95.60 | 0.25 (0.03% relative), 0.07 (0.07% relative) | 2,000/bottle | 15 mL | 12011-20249 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2571 | Ethanol standard solution (70%ABV) (TRM-S-5049b) | Density (in air) at 20 oC (kg/m3) = 885.32, Alcoholic Strength (% ABV) = 70.08 | 0.29 (0.03% relative), 0.05 (0.07% relative) | 2,000/bottle | 15 mL | 12011-20250 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2572 | Ethanol standard solution (40%ABV) (TRM-S-5049c) | Density (in air) at 20 oC (kg/m3) = 948.01, Alcoholic Strength (% ABV) = 40.04 | 0.33 (0.03% relative), 0.16 (0.4% relative) | 2,000/bottle | 15 mL | 12011-20251 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2573 | Ethanol standard solution (15%ABV) (TRM-S-5049d) | Density (in air) at 20 oC (kg/m3) = 978.94, Alcoholic Strength (% ABV) = 15.03 | 0.28 (0.03% relative), 0.02 (0.13% relative) | 2,000/bottle | 15 mL | 12011-20252 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
149 | High purity alpha-mangostin standard | 920.7 mg/g | 0.8 %relative | 15,000/bottle | 50 mg | 12011-20254 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
150 | High purity mitragynine standard | 980.4 mg/g | 0.4 %relative | 10,000/bottle | 10 mg | 12011-20255 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
151 | High purity [6]-gingerol standard | 972.8 mg/g | 0.9 %relative | 5,000/bottle | 25 mg | 12011-20256 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2549 | Caffeine standard solution for other customer/ made to order | TRM-S-5054 | 0.5 mg/mL-200 mg/mL | 5,000/set (1 set = 8 ampoule) | 1 mL/ampoule | 12011-20258 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2567 | High purity terpinen-4-ol standard (TRM-S-5055) | 98.70% | 7.6%relative | 5,000 baht/bottle | 50 mg | 12011-20259 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2574 | High purity curcumin standard (TRM-S-5056) | 993.1 mg/g | 3.5% relative | 5,000 baht/bottle | 50 mg | 12011-20260 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2575 | High purity demethoxycurcumin standard (TRM-S-5057) | 982.8 mg/g | 3.5% relative | 5,000 baht/bottle | 20 mg | 12011-20261 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2568 | Sodium Bromide Standard Solution(Density) (TRM-S-5059) | 600 to 1700 mg/g | 0.01-1.0% | 2,000 baht/ampoule | 15 mL | 12011-20263 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
238 | Glucose in frozen human serum | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate |  | 5,000/set | 3 levels, 1mL/vial | 12011-20502 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
181 | Total malachite green in shrimp | 18.35 µg/kg | 4.00 % relative | 5,000/ bottle |  | 12011-20601 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
182 | Aflatoxins in peanut butter | 2-10 ng/g | 8-10% relative | 4,500 /bottle | 30 g/bottle | 12011-20602 | Chemical Metrology Department | 4500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
183 | Melamine in cracker | 1.56 mg/kg | 9% relative | 3,000/bottle |  | 12011-20603 | Chemical Metrology Department | 3000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
186 | Melamine in Milk Powder | 1.03 mg/kg | 16% relative | 3,000/20g bottle |  | 12011-20604 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
187 | Fatty Acid in Cooking Oil | Palmitic acid 106 mg/g,11.4% Stearic acid 33 mg/g, 4.1% Oleic acid 231 mg/g, 25.7% Linoleic acid 493 mg/g, 53.4% Linolenic acid 46 mg/g, 5.3% | 7-10% relative | 6000/bottle |  | 12011-20605 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
189 | Aflatoxin in rice flour | AFB1: 11.5 ng/g AFB2: 8.7ng/g AFG1: 7.5 ng/g AFG2: 3.0 ng/g | AFB1: 1.5 ng/g AFB2: 0.8 ng/g AFG1: 1.0 ng/g AFG2: 0.4 ng/g | 4,500baht/pouch | 80g/pouch | 12011-20609 | Chemical Metrology Department | 4500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
191 | Andrographolide in green chiretta | 18.7 mg/g | 7% relative | 4,000 baht/bottle |  | 12011-20611 | Chemical Metrology Department | 4000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
192 | THC and CBD in hemp oil standard (CRM) | THC: 1.47 mg/g CBD: 53 mg/g | THC:30 %relative CBD: 17 %relative | 8,000/bottle | 5 ml | 12011-20614 | Chemical Metrology Department | 8000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
225 | Organochlorine Pesticides in Soil | Endosulfan II 0.43 mg/kg Endosulfan sulfat 0.56 mg/kg | 14% Relative 15% Relative | 3,000/bottle |  | 12011-20801 | Chemical Metrology Department | 3000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2566 | % Fiber from sugar cane (TRM-M-5001) | 11.62% MC | 0.6% MC | 5000 baht/bag | 50 g/bag | 12011-20901 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
184 | Clenbuterol in feed | 71.9 µg/kg | 5.43 % relative | 2,850/bottle |  | 12011-20901 | Chemical Metrology Department | 2850 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
185 | Salbutamol in feed | 79.5 µg/kg | 14.21 % relative | 2,850/bottle |  | 12011-20902 | Chemical Metrology Department | 2850 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
252 | Determination of organic componds in environment | ng/g - mg/g | 3-15% relative | 10,000-20,000/sample | 1. Determination of BTEX in water | 12011-60210 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
188 | Organophosphates pesticides in rice flour | Diazinon 15 ng/g Chlopyrifos 18 ng/g | 6-8% relative | 4000/bottle | 30g/bottle | 12011-60606 | Chemical Metrology Department | 4000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
102 | 5% Brix Sucrose Standard Solution | 5.00 % Brix, 1.34026 nD | 0.01 % Brix, 0.00007 nD | 3,600/15 mL |  | 12012-20201 | Chemical Metrology Department | 3600 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
245 | Determination of purity of organic compounds (Mass balance method) | as specification | as specification | 10,000-19,000.00/sample | 1. Purity assessment of ethanol 2. Purity assessment of mitragynine 3. Purity assessment of 6-gingerol 4. Purity assessment of mangostin 5. Purity assessment of curcumin 6. Purity assessment of demethoxycurcumin | 12012-60101 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2540 | Determination of purity of organic compounds (Mass balance method) | as specification | as specification | 10,000-19,000.00/sample | 1. Purity assessment of ethanol 2. Purity assessment of mitragynine 3. Purity assessment of 6-gingerol 4. Purity assessment of mangostin 5. Purity assessment of curcumin 6. Purity assessment of demethoxycurcumin 7. Purity assessment of bisdemethoxycurcumin | 12012-60101 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
246 | Determination of purity of organic compounds (specified technique) | as specicication | as specification | 2,000-10,000.00/sample | 1. Thermo-physical properties by DSC | 12012-60102 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
248 | Determination of Brix and refractive indices in food and beverages by refractometer | 5 - 60 %Brix 1.34026 - 1.44193 nD | 0.04 %Brix 0.00008 nD | 1,500/sample |  | 12012-60201 | Chemical Metrology Department | 1500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
249 | Determination of polarization in food and beverages by polarimeter | 0 - 100 oZ 0 - 40 OR | 0.3 - 0.6 oZ 1 - 3 OR | 11,000/sample |  | 12012-60202 | Chemical Metrology Department | 11000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
250 | Determination of Aflatoxins in solution by LC | ng/g - mg/g | 3 - 10 %relative | 6,000/analyte +3,000 for next analyte |  | 12012-60203 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
251 | Determination of sodium chloride by refractometry | 0-25 g/100g | 1.72 % relative | 6,000/sample | calibration curve | 12012-60204 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
253 | Determination of amylose in rice by size exclusion chromatography (SEC) technique | 0-30% | as specification | 6,000/sample |  | 12012-60302 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
254 | Determination of amylose in rice by Iodine binding assay using a plate reader | 0-30% | as specification | 6,000/sample |  | 12012-60303 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
255 | Determination of dietary suppliment adulteration in food | 100-1000 ng/g | 7-8 % relative | 18,000/sample | Exact-matching IDMS | 12012-60304 | Chemical Metrology Department | 18000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
242 | Determination of organic compounds in biological fluid by IDMS | ug/g - mg/g | 3-20% relative | 10,000 - 15,000/sample | 1. Determination of glucose by IFCC reference method 2. Determination of peptide by amino acid hydrolysis 3. lipoprotein | 12012-60501 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
243 | Determination of organic compounds in food and feed by IDMS | ng/g - mg/g | 0.1 - 15 %relative | 10,000-20,000/sample | 1. Determination of Ametryn, Atrazine, and Pendimethalin in fruit and vegetable 2.Determination of Metalaxyl, Thiabendazole, Prochloraz, Tebuconazole, Difenoconazole and Azoxystrobin in fruit and vegetable 3. metronidazole 4. dimetridazole 5. amino acids | 12012-60601 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
244 | Determination of organic compounds in food and feed by calibration curve | ng/g - mg/g |  | 10,000 /analyte + 5,000 for next analyte | 1. Determination of nitrosamines in food 2. Determination of B vitamins in food 3. Determination of Ametryn, Atrazine, and Pendimethalin in fruit and vegetable 4. Determination of Metalaxyl, Thiabendazole, Prochloraz, Tebuconazole, Difenoconazole and Azoxystrobin in fruit and vegetable | 12012-60602 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2541 | Determination of organic compounds in food and feed by calibration curve | ng/g - mg/g |  | 10,000 /analyte + 5,000 for next analyte | 1. Determination of nitrosamines in food 2. Determination of B vitamins in food 3. Determination of Ametryn, Atrazine, and Pendimethalin in fruit and vegetable 4. Determination of Metalaxyl, Thiabendazole, Prochloraz, Tebuconazole, Difenoconazole and Azoxystrobin in fruit and vegetable | 12012-60602 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2545 | Determination of organic compounds in botanical products | ng/g-mg/g | 5-20% relative | 5,000-15,000/sample | This service includes; 1. Determination of Gamma-oryzanol in rice 2.Determination of anthocyanins in color rice 3. Determination of active ingredient in herb 4. ochratoxin 5. ochratoxin 6. pesticides 7. mitragynine | 12012-60901 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
247 | Determination of organic compounds in botanical products | ng/g-mg/g | 5-20% relative | 10,000-15,000/sample | This service includes; 1. Determination of Gamma-oryzanol in rice 2.Determination of anthocyanins in color rice 3. Determination of active ingredient in herb 4. ochratoxin 5. ochratoxin 6. pesticides 7. mitragynine | 12012-60901 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
174 | Trace and Essential Elements in Prawn »ÃÔÁÒ³¸ÒµØã¹¡Øé§ | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate |  | 2,000/ bottle | 5 g/bottle | 12022-20002 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
226 | Elements in Acrylonitrile Butadiene Styrene (ABS) Plastic, Low Levels »ÃÔÁÒ³¸ÒµØÃÐ´ÑºµèÓã¹¾ÅÒÊµÔ¡ª¹Ô´ ABS | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕ´Âã¹µÑÇÍÂèÒ§ãº Certificate |  | 5,000/ bottle | 30 g/bottle | 12022-20003 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
229 | Elements in Human Serum »ÃÔÁÒ³¸ÒµØã¹«ÕÃÑèÁ | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕ´Âã¹µÑÇÍÂèÒ§ãº Certificate |  | 1,000/bottle | 4 mL/bottle | 12022-20006 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
193 | Trace Elements in Water »ÃÔÁÒ³¸ÒµØã¹¹éÓ | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕ´Âã¹µÑÇÍÂèÒ§ãº Certificate |  | 2,000/bottle | 100 mL/bottle | 12022-20008 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
194 | Mercury in Water »ÃÔÁÒ³»ÃÍ·ã¹¹éÓ | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕ´Âã¹µÑÇÍÂèÒ§ãº Certificate |  | 1,700/bottle | 10 mL/bottle | 12022-20009 | Chemical Metrology Department | 1700 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
176 | Elements, AB and MeHg in Tuna Tissue | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate |  | 5,000/bottle | 20 g/bottle | 12022-20010 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
177 | Elements in Drinking Water | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate |  | 5,500/bottle | 100 ml/bottle | 12022-20011 | Chemical Metrology Department | 5500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
178 | Elements in Green Chiretta (Andrographis Paniculata) Powder | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate |  | 6,000/bottle | 25 g/bottle | 12022-20012 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
227 | Elements in 10-mL Digesting Solution from Cotton Swab | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕ´Âã¹µÑÇÍÂèÒ§ãº Certificate |  | 4,000/package | 10 cotton swabs/ package | 12022-20013 | Chemical Metrology Department | 4000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
179 | Elements in Turmeric powder | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate | 6,000/bottle | 30 g/bottle | 12022-20014 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
228 | Toxic Elements in Lipstick | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕ´Âã¹µÑÇÍÂèÒ§ãº Certificate |  | 3,500/bottle | 10 g/bottle | 12022-20015 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
180 | Elements in Animal Feed | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate | 3,000/bottle | 10 g/bottle | 12022-20016 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
230 | Electrolytes in human serum | Na: 3,197.99 mg/kg K: 151.97 mg/kg Ca: 89.21 mg/kg Cl: 3,718.96 mg/kg | (2 - 3)% relative | 2,800/4 mL |  | 12022-20017 | Chemical Metrology Department | 2800 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2548 | Toxic Elements in Cosmetic cream containing hemp-derived ingredients | TRM-M-2005 |  | 2,500/bottle | 5 g/bottle | 12022-20018 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
175 | Elements and Arsenic Species in White Rice Flour | see certificate â»Ã´´ÙÃÒÂÅÐàÍÕÂ´ã¹µÑÇÍÂèÒ§ãº Certificate |  | 4,000/bottle | 25 g/bottle | 12023-20007 | Chemical Metrology Department | 4000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
264 | Determination of Elements in sample by ID-ICP/MS | 0.0001-30000 mg/kg | 0.7-7.5%relative | 20,000 for first element + 3000 for next element |  | 12023-60101 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2537 | Determination of Elements in sample by ID-ICP/MS | 0.0001-30000 mg/kg | 0.3-7.5%relative | 20,000 for first element + 3000 for next element |  | 12023-60101 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
265 | Determination of Elements in sample by gravimetric standard addition | 0.0004-300000 mg/kg | 1.5-18 %relative | 17,500 for first element + 3000 for next element |  | 12023-60102 | Chemical Metrology Department | 17500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2538 | Determination of Elements in sample by gravimetric standard addition | 0.0004-300000 mg/kg | 1.5-18 %relative | 17,500 for first element + 3000 for next element |  | 12023-60102 | Chemical Metrology Department | 17500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
266 | Determination of Elements in sample by external calibration | 0.000002 - 300000 mg/kg | 1.0-16%relative | 10,000 for first element + 3000 for next element |  | 12023-60103 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2536 | Determination of Elements in sample by external calibration | 0.000002 - 300000 mg/kg | 1.0-16%relative | 10,000 for first element + 3000 for next element |  | 12023-60103 | Chemical Metrology Department | 10000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
267 | Determination of Elemental speciation in sample | 0.02 - 3200 mg/kg | 3.5-10%relative | 17,500 for first specy + 3000 for next species |  | 12023-60104 | Chemical Metrology Department | 17500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2547 | Determination of Particle number concentration in sample by spICP-MS | (0.5-2.0)*E11 particles/g | 13.2 -18%relative | 15,000 /sample |  | 12023-60105 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2539 | Determination of lead isotope ratio in fresh water by single collector ICP-MS | 2.5-50 ng/g | 0.03 - 0.5% relative | 20,000 / sample |  | 12023-60106 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
5 | pH meter with associated electrode | 1 - 12 pH | 0.01 pH | 4,000/set (3 point + 500 next point) |  | 12031-10101 | Chemical Metrology Department | 4000 | Chemical Metrology Department | 7 | meter, power, phase, ac, wattmeter
54 | Primary Phthalate buffer solution | 3.990 - 4.020 pH | 0.005 pH | 15,000/500 mL |  | 12031-20109 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
55 | Primary Phosphate buffer solution | 6.85 - 6.88 pH | 0.005 pH | 15,000/500 mL |  | 12031-20110 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
56 | Primary Borate buffer solution | 9.16 - 9.20 pH | 0.005 pH | 15,000/500 mL |  | 12031-20111 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
60 | Primary Carbonate buffer solution | 9.990 - 10.020 pH | 0.006 pH | 15,000/500 mL |  | 12031-20120 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
57 | Working pH standard (Phthalate) | 3.99 - 4.02 pH | 0.01 pH | 1,000/500 mL |  | 12031-20121 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
58 | Working pH standard (Equipmolal phosphate) | 6.85 - 6.88 pH | 0.01 pH | 1,000/500 mL |  | 12031-20122 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
59 | Working pH standard (Tetraborate) | 9.17 - 9.20 pH | 0.01 pH | 1,000/500 mL |  | 12031-20123 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
61 | Working pH standard (Carbonate) | 9.99 - 10.02 pH | 0.01 pH | 1,000/500 mL |  | 12031-20124 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
64 | Working pH standard (phosphate) | 6.99 - 7.03 pH | 0.01 pH | 1,000/500 mL |  | 12031-20125 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
62 | Working pH standard (0.01 N Hydrochloric acid) | 1.08 - 1.12 pH | 0.02 pH | 1,000/500 mL |  | 12031-20126 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
63 | Working pH standard (Tetroxalate) | 1.66 - 1.70 pH | 0.01 pH | 1,000/500 mL |  | 12031-20127 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
65 | Working pH standard (alkaline phosphate) | 11.70 - 11.74 pH | 0.01 pH | 1,000/500 mL |  | 12031-20128 | Chemical Metrology Department | 1000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2565 | pH in wastewater (TRM-E-2005) | pH 6.50-7.50 | 0.04pH | 300/ bottle | 100 mL/bottle | 12031-20129 | Chemical Metrology Department | 300 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
70 | Potassium Dichromate Solution 5 concentration | 20, 40, 60, 80 and 100 mg/kg (each set choose from the concentrations above) | 0.01 A | 43,000/set |  | 12031-20201 | Chemical Metrology Department | 43000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
69 | Potassium Dichromate Solution 4 concentrations | 20, 40, 60, 80 and 100 mg/kg (each set choose from the concentrations above) | 0.01 A | 38,000/set |  | 12031-20202 | Chemical Metrology Department | 38000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
68 | Potassium Dichromate Solution 3 concentrations | 20, 40, 60, 80 and 100 mg/kg (each set choose from the concentrations above) | 0.01 A | 32,800/set |  | 12031-20203 | Chemical Metrology Department | 32800 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
67 | Potassium Dichromate Solution 2 concentrations | 20, 40, 60, 80 and 100 mg/kg (each set choose from the concentrations above) | 0.01 A | 27,700/set |  | 12031-20204 | Chemical Metrology Department | 27700 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
66 | Potassium Dichromate Solution 1 concentration | 20, 40, 60, 80 and 100 mg/kg (each set choose from the concentrations above) | 0.01 A | 22,600/set |  | 12031-20205 | Chemical Metrology Department | 22600 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
71 | Potassium Iodide Standard Solution | 10 g/L | 0.5 nm | 20,000/set |  | 12031-20206 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
72 | Zinc standard solution | 10,000 ppm | 1% relative | 1,250/100 ml |  | 12031-20301 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
73 | Cadmium standard solution | 10,000 ppm | 1% relative | 1,250/100 ml |  | 12031-20302 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
76 | Arsenic standard solution | 1000 ppm | 0.2% rlative | 2,000/100 ml |  | 12031-20303 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
74 | Chloride standard solution | 100 ppm | 1% relative | 1,250/100 ml |  | 12031-20401 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
97 | Nitrate standard solution | 100 mg/kg | 1% relative | 1,500 / bottle |  | 12031-20402 | Chemical Metrology Department | 1500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
98 | Fluoride standard solution | 1000 mg/kg | 1% relative | 1,500 / bottle |  | 12031-20403 | Chemical Metrology Department | 1500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
99 | sulfate standard solution | 1000 mg/kg | 3%relative | 1,250/100 mL |  | 12031-20404 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
100 | Bromide standard solution | 900-1,200 mg/kg | 5% relative | 1,250/125 mL |  | 12031-20405 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
75 | Sodium standard solution | 100 ppm | 1% relative | 1,250/100 ml |  | 12031-20501 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
101 | Magnesium standard solution | 900-1,200 mg/kg | 3% relative | 1,250/125 mL |  | 12031-20502 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2577 | Calcium standard solution (TRM-S-2057) | 900-1,200 mg/kg | 4%relative | 1,250/ bottle | 100 mL/bottle | 12031-20503 | Chemical Metrology Department | 1250 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
80 | Secondary electrolytic conductivity solution (0.1 mol/l KCl) | 12.88 mS/cm | 0.5% relative | 5,000/250 ml |  | 12031-20601 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
81 | Secondary electrolytic conductivity solution (0.01mol/l KCl) | 1413 mS/cm | 0.5% relative | 5,000/250 ml |  | 12031-20602 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
84 | Secondary electrolytic conductivity standard solution (nominall 50 mS/cm) | 50 mS/cm | 1% relative | 5,000/250 ml |  | 12031-20605 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
89 | Working electrolytic conductivity solution (0.1 mol/l KCl) | 12.88 mS/cm | 1.5 %relative | 3,000/250 ml |  | 12031-20606 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
90 | Working electrolytic conductivity solution (0.01 mol/l KCl) | 1413 mS/cm | 1.5 %relative | 3,000/250 ml |  | 12031-20607 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
85 | Secondary electrolytic conductivity solution (nominal 25 µS/cm) | 25 µS/cm | 1% relative | 5,000 /250 ml |  | 12031-20608 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
91 | working electrolytic conductivity standard solution (0.001 mol/l KCl) | 147 µS/cm | 1.5% relative | 2,500/ 250 ml |  | 12031-20609 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
92 | working electrolytic conductivity standard solution (Nominal 84 µS/cm) | 84 µS/cm | 2% relative | 2,500/ 250 ml |  | 12031-20610 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
93 | Working electrolytic conductivity standard solution (Nominal 50 µS/cm) | 50 µS/cm | 2% relative | 2,500/ 250 ml |  | 12031-20611 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
86 | Secondary electrolytic conductivity solution (nominal 10 mS/cm) | 10 mS/cm | 2.2 % relative | 2,500/125 ml |  | 12031-20612 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
87 | Secondary electrolytic conductivity solution (1 mol/L KCl) | 111.3 mS/cm | 0.5 %relative | 5,000/250 ml |  | 12031-20613 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
94 | Working electrolytic conductivity standard solution (nominal 25 µS/cm) | 25 mS/cm | 2% relative | 2,500/250 ml |  | 12031-20614 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
95 | Secondary electrolytic conductivity solution (Nominal 5 µS/cm) | 5 µS/cm | 2.5% relative | 2,500/250 ml |  | 12031-20615 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
88 | Working electrolytic conductivity standard solution (1 mol/L KCl) | 111.3 mS/cm | 0.7% relative | 3,000/250 mL |  | 12031-20616 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2550 | Working electrolytic conductivity standard solution (1 mol/L KCl) | TRM-S-2053 | 0.7% relative | 3,000/250 mL |  | 12031-20616 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
96 | Working electrolytic conductivity standard solution (Nominal 10 uS/cm ) | 10 mS/cm | 5% relative | 1,300/125 mL |  | 12031-20617 | Chemical Metrology Department | 1300 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
2569 | Working electrolytic conductivity standard solution (Nominal 10 uS/cm) (TRM-S-2054) | 10 mS/cm | 5% relative | 1,300/125 mL |  | 12031-20617 | Chemical Metrology Department | 1300 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
77 | pH of ethanol | 6.0 - 8.0 pH | 0.5 pH | 2,500 / bottle |  | 12031-20701 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
78 | electrolytic conductivity of ethanol | 0.1-1 uS/cm | 5%relative | 2,500/bottle | capacity 50 ml | 12031-20702 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
79 | Water content in Ethanol | 0.1 - 1.0 mg/g | 5% relative | 3,000/5 ampoule |  | 12031-20703 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
195 | Chloride in Ethanol | 1.51 mg/kg | 6% relative | 1,800/250 mL | 250 mL | 12031-20704 | Chemical Metrology Department | 1800 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
152 | Platinum electrode for pH measurement |  |  |  | For differential cell and Harned cell method | 12031-20801 | Chemical Metrology Department | 7880 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
239 | 65% Nitric acid, suprapure grade, 500 ml | - | - | 3,500/ bottle | 500 ml / bottle | 12031-20803 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
240 | 65% Nitric acid, suprapure grade, 1000 ml | - | - | 6,000/bottle | 1,000 ml / bottle | 12031-20804 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
241 | Tune solution for iCAP | 1 ng/g | - | 5,000/bottle | 500 ml / bottle | 12031-20805 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
262 | Determination of substances in sample (Electrochemical method (cyclic voltammetry) | (5 - 100 )mg/kg | 3 % relative | 3,000 / sample | 1. nitrite 2. NaCl | 12031-60103 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
263 | Measurement of acidity in ethanol | (5 - 100 )mg/kg | 3 % relative | 3,000 / sample |  | 12031-60401 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
82 | Secondary electrolytic conductivity solution (0.001mol/l KCl) | 147 mS/cm | 0.8% relative | 5,000/250 ml |  | 12031–20603 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
83 | Secondary electrolytic conductivity solutionn (Nominal 84 µS/cm) | 84 mS/cm | 1% relative | 5,000/250 ml |  | 12031–20604 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
1 | Piston pipette | 20 mL to 10 mL | 0.04 mL | 4,000 three point + 500 next point |  | 12032-10201 | Chemical Metrology Department | 4000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2 | Volumetric apparatus | 1 mL to 2000 mL | Class A | 2,500 two point + 500 next point | volumetric flask, pipette, burette, pycnometer | 12032-10203 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
3 | recalibrate of referene material for calibration UV/VIS spectrophotometer (Photometric accuracy) | 20 to 100 mg/kg | 0.003 A | 7,500/set | potassium dichromate solution | 12032-10301 | Chemical Metrology Department | 7500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
4 | recalibrate of referene material for calibration UV/VIS spectrophotometer (stray light) | 190 - 280 nm | 0.12 nm | 6,500/set |  | 12032-10306 | Chemical Metrology Department | 6500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
46 | Secondary pH standard (Phthalate) | 3.990 - 4.020 pH | 0.006 pH | 8,000/set (5 x 100 mL) | Differential cell method | 12032-20101 | Chemical Metrology Department | 8000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
48 | Secondary pH standard (Equipmolal phosphate) | 6.850 - 6.880 pH | 0.006 pH | 8,000/set (5 x 100 mL) | Differential cell method | 12032-20102 | Chemical Metrology Department | 8000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
50 | Secondary pH standard (Tetraborate) | 9.17 - 9.20 pH | 0.007 pH | 8,000/set (5 x 100 mL) | Differential cell method | 12032-20103 | Chemical Metrology Department | 8000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
52 | Secondary pH standard (Carbonate) | 9.99 - 10.02 pH | 0.006 pH | 8,000/set (5 x 100 mL) | Differential cell method | 12032-20104 | Chemical Metrology Department | 8000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
47 | Secondary pH standard (Phthalate) | 3.990 - 4.020 pH | 0.006 pH | 1,700/100 mL | Differential cell method | 12032-20105 | Chemical Metrology Department | 1700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
49 | Secondary pH standard (Equipmolal phosphate) | 6.850 - 6.880 pH | 0.006 pH | 1,700/100 mL | Differential cell method | 12032-20106 | Chemical Metrology Department | 1700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
51 | Secondary pH standard (Tetraborate) | 9.17 - 9.20 pH | 0.007 pH | 1,700/100 mL | Differential cell method | 12032-20107 | Chemical Metrology Department | 1700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
53 | Secondary pH standard (Carbonate) | 9.99 - 10.02 pH | 0.006 pH | 1,700/100 mL | Differential cell method | 12032-20108 | Chemical Metrology Department | 1700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
44 | Secondary pH standard (tetroxalate) | 1.66-1.70 pH | 0.006 pH | 8,000/set (5 x 100 mL) | Differential cell method | 12032-20109 | Chemical Metrology Department | 8000 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
45 | Secondary pH standard (tetroxalate) | 1.66-1.70 pH | 0.006 pH | 1,700/100 mL | Differential cell method | 12032-20110 | Chemical Metrology Department | 1700 | Chemical Metrology Department | 4 | standard, solution, mass, set, ph
257 | Assigned pH value for sample (secondary method) | 1 - 12 pH | 0.004 pH | 13,000/sample |  | 12033 - 60109 | Chemical Metrology Department | 13000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
261 | Determination of cation and anion in sample (ion chromtography) | as specification | %1 - 2%relative | 7,000/sample |  | 12033-60103 | Chemical Metrology Department | 7000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
258 | Assigned pH value for sample (working method) | 1 - 12 pH | 0.01 pH | 2,500/sample |  | 12033-60106 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
256 | Assigned pH value for sample (primary method) | 1 - 12 pH | 0.003 pH | 13,750/sample |  | 12033-60107 | Chemical Metrology Department | 13750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
259 | Assigned electrolytic conductivity value for sample (secondary method) | 5 mS/cm - 12.88 mS/cm | 0.3% - 1% relative | 11,000 /sample |  | 12033-60404 | Chemical Metrology Department | 11000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
260 | Assigned electrolytic conductivity value for sample (commercial cell) | 84 - 12880 mS/cm | 0.8% - 1.3% relative | 2,500/sample |  | 12033-60406 | Chemical Metrology Department | 2500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
24 | MON810 plasmid DNA | 1.04 copy number ratio | 0.061 copy number ratio | 12500 |  | 12042-20101 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
153 | Pure pork meat | 1000 g/kg |  | 3500 |  | 12042-20102 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
154 | Pure beef meat | >998.8 g/kg |  | 3500 |  | 12042-20103 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
23 | Pork DNA solution | 1.2 x 105 copies/µL | 1.3 x 104 copies/µL | 5000 |  | 12042-20104 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
25 | TRM-enzyme a-amylase | 382.8-457.2 U/L | 37.2 U/L | 12500 |  | 12042-20105 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
155 | Blank soybean GTS 40-3-2 matrix | <0.2 g/kg |  | 3500 |  | 12042-20106 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
26 | Beef DNA solution | 1.0 x 105 copies/µL | 1.3 x 104 copies/µL | 5000 |  | 12042-20107 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
156 | 1% Pork in beef matrix | 10 g/kg | 2.20 g/kg | 3500 |  | 12042-20108 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
157 | 10% Pork in beef matrix | 100 g/kg | 11 g/kg | 3500 |  | 12042-20109 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
27 | Blank papaya DNA solution | 1.8 x 104 copies/µL | 2.3 x 103 copies/µL | 5000 |  | 12042-20110 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
158 | Blank meatball matrix | 1 000g/kg |  | 3500 |  | 12042-20111 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
159 | 0.1% Pork in beef matrix | 1 g/kg | 0.33 g/kg | 3500 |  | 12042-20112 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
160 | 1% beef in feed matrix | 10 g/kg | 0.1 g/kg | 3500 |  | 12042-20113 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
28 | Roundup ready plasmid DNA | 1.09 copy number ratio | 0.18 copy number ratio | 12,500 |  | 12042-20114 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
29 | Bt63-Rice plasmid DNA | 2.3 x 104 copies/µL | 4.7 x 103 copies/µL | 12,500 |  | 12042-20115 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
30 | Papaya line 55-1 plasmid DNA | 8.8 x 10^5 copies/µL | 1.8 x10^5 copies/µL | 12500 |  | 12042-20116 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
162 | 0.1% (w/w) beef in feed | 1 g/kg | 0.15 g/kg | 3500 |  | 12042-20117 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
161 | 1 % Pork in meatball matrix | 10.1 g/kg | 1.1 g/kg | 3500 |  | 12042-20118 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
31 | Beef DNA Solution (Processed food) | 1.3 x 105 copies/µL | 1.4 x 104 copies/µL | 5000 |  | 12042-20119 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
32 | %Pork in beef DNA solution (Processed food) | 1.37% cp/cp | 0.19% cp/cp | 5,000 |  | 12042-20120 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
163 | Pure pork meat (processed food) | 1,000 g/kg |  | 3,500/sample |  | 12042-20121 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
164 | Pure Beef meat (processed food) | 1,000 g/kg |  | 3,500/sample |  | 12042-20122 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
165 | 1% (w/w) pork in beef matrix (process food) | 10g/kg | 2g/kg | 3,500/sample |  | 12042-20123 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
166 | 7% (w/w) pork in beef matrix (process food) | 70g/kg | 5g/kg | 3,500/sample |  | 12042-20124 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
33 | LLrice62 Rice-Plasmid DNA | 1.2x 105 copies/µL | 2.8 x 104 copies/µL | 12500 |  | 12042-20124 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
34 | Pork DNA solution (processed food) | 1x105 copies/µL | 1x104 copies/µL | 5000 |  | 12042-20125 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
35 | ASFV in tissue homogenate | 358 copies/µL | 128 copies/µL | 12500 |  | 12042-20126 | Chemical Metrology Department | 12500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
36 | Set of ASFV in tissue homogenate | 5-1500 copies/µL |  | 12,500/set |  | 12042-20127 | Chemical Metrology Department | 12500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
37 | MON89788 plasmid DNA | 2.5 x 105 copies/µL | 4.7 x 104 copies/µL | 12500 |  | 12042-20128 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
38 | Goat DNA Solution | 8.8 ×104 copies/µL | 1.4 ×104 copies/µL | 5,000 |  | 12042-20129 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
39 | LLrice601 Rice-Plasmid DNA | 1.2x 105 copies/µL | 2.6 x 104 copies/µL | 12500 |  | 12042-20130 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
167 | Pure chicken meat | 1,000 g/kg |  | 3500 |  | 12042-20131 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
168 | Pure Goat meat | 1,000 g/kg |  | 3500 |  | 12042-20132 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
169 | Curcuminoids in turmeric powder: Curcumin DMC BDMC | 30.2 mg/g 11.8 mg/g 11.6 mg/g | 6.4 mg/g 0.6 mg/g 2.0 mg/g | 4000 |  | 12042-20133 | Chemical Metrology Department | 4000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
40 | Chicken DNA Solution | 2.5 × 105 copies/µL | 3.1 × 104 copies/µL | 5,000 |  | 12042-20134 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
41 | Papaya Huanong1 Plasmid DNA | 1.9 × 105 copies/µL | 2.0 ×104 copies/µL | 12,500 |  | 12042-20135 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
42 | CaMV35S Plasmid DNA | 1.2 × 104 copies/µL | 1.8 × 103 copies/µL | 12,500 |  | 12042-20136 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
170 | Pure sheep meat | 1,000 g/kg |  | 3500 |  | 12042-20137 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
171 | 0.1% (w/w) Pork in plant based | 1 g/kg | 0.05 g/kg | 3500 |  | 12042-20138 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
172 | 0.1% (w/w) Beef in plant based | 1 g/kg | 0.06 g/kg | 3500 |  | 12042-20139 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
173 | 0.1% (w/w) Chicken in plant based | 1 g/kg | 0.06 g/kg | 3500 |  | 12042-20140 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
43 | Sheep DNA solution | 1.1 × 105 copies/µL | 1.7 × 104 copies/µL | 5,000 |  | 12042-20141 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
2560 | Fish DNA Solution (TRM S-1021) | 1.9 x 105 ?L-1 | 3.3 x 104 ?L-1 | 5000 |  | 12042-20143 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
2561 | Pure crocodile meat | 1,000 g/kg |  | 3,500 |  | 12042-20144 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2562 | Crocodile DNA solution | 1.2 x 105 ?L-1 | 2.1 x 104 ?L-1 | 5,000 |  | 12042-20145 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
2563 | Pure fish meat | 1,000 g/kg |  | 3,500 |  | 12042-20146 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2576 | 0.1% (w/w) Goat in plant based (TRM F-1025) | 1 g/kg | 0.07 g/kg | 3,500 |  | 12042-20147 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2564 | p72-ASFV plasmid DNA | 1.6 ? 105 ?L-1 | 3.9 ? 104 ?L-1 | 12,500 |  | 12042-20148 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
268 | Catalytic activity concentration of enzyme by microplate reader | U/L | U/L | 6,000/sample |  | 12043-60101 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
269 | Qualification/quantification of DNA concentration in Foodstuffs/Feed/Biomarkers by droplet digital Polymerase Chain Reaction (ddPCR) | detection or not detection/ copy number or ratio |  | 12,500/sample | Quantification of LLrice62 by duplex ddPCR LOQ - LLrice 62: 8.41±1.11 copies/µL - SPS: 8.49± 1.30copies/µL Quantification of LLrice601 by duplex ddPCR LOQ - LLrice601: 7.77±1.21 copies/µL - SPS: 7.51± 0.99copies/µL HIV-1 RNA copy number quantification - 145-757 copies/µL Detection of chicken DNA by ddPCR LOD: 4.49 ± 0.96 cp/µL Detection of ASFV by ddPCR LOD: 1.89 ± 0.55 cp/µL | 12043-60102 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
2543 | Qualification/quantification of DNA concentration in Foodstuffs/ Feed/ Biomarkers by droplet digital Polymerase Chain Reaction (ddPCR) | detection or not detection/ copy number or ratio |  | 12,500/sample | Quantification of LLrice62 by duplex ddPCR LOQ - LLrice 62: 8.41?1.11 copies/?L - SPS: 8.49? 1.30copies/?L Quantification of LLrice601 by duplex ddPCR LOQ - LLrice601: 7.77?1.21 copies/?L - SPS: 7.51? 0.99copies/?L HIV-1 RNA copy number quantification - 145-757 copies/?L Detection of chicken DNA by ddPCR LOD: 4.49 ? 0.96 cp/?L Detection of ASFV by ddPCR LOD: 1.89 ? 0.55 cp/?L | 12043-60102 | Chemical Metrology Department | 12500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
270 | DNA Quantification in Foodstuffs and feed by quantitative real-time PCR (qPCR) | copy number or ratio |  | 8,500/sample |  | 12043-60103 | Chemical Metrology Department | 8500 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
271 | DNA Qualification in Foodstuffs/Feed/Biomarkers quantitative by real-time PCR (qPCR) | detection or not detection |  | 5,000/sample | - The detection of Mycoplasma arginini DNA using Real-Time PCR - The detection of Mycoplasma fermentans DNA using Real-Time PCR - Detection of chicken DNA by qPCR | 12043-60104 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
2546 | Detection of pork DNA using Real-Time PCR | Quality measurement | Quality measurement | 5000 |  | 12043-60104 | Chemical Metrology Department | 5000 | Chemical Metrology Department | 5 | digital, dna, analog, plasmid, dial
272 | Identification of biomolecules by MALDI-TOF MS | Quality measurement | Quality measurement | 2,000/sample |  | 12043-60105 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2544 | Identification of biomolecules by MALDI-TOF MS | Quality measurement | Quality measurement | 2,000/sample |  | 12043-60105 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
273 | qualification of protein in Foodstuffs and feed by LC Q-TOF MS | detection or not detection |  | 3,500/sample |  | 12043-60106 | Chemical Metrology Department | 3500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
274 | Confocal Laser Scanning Microscopes ZEISS LSM880 | Cell imaging | Cell imaging | 2,500/hr |  | 12043-60107 | Chemical Metrology Department | 2500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
275 | Nanoparticle size measurement in solution | 10 - 1000 nm |  | 2,000/sample |  | 12043-60108 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
276 | Identification and quantification of organic compounds in herb by GC-MS and HPLC | Quality measurement / Quantity measurement | Quality measurement / Quantity measurement | 2,000/sample | Quantification of terpinen-4-ol in Plai extract using GC/MS - range 5-100 ug/g - uncertainty 7-8% relative Quantification of curcumin MS in turmeric extract using HPLC-DAD - range 4-300 ug/g - uncertainty 7-8% relative | 12043-60109 | Chemical Metrology Department | 2000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
277 | Detection of Foodstuffs/Feed/Biomarkers using loop-mediated isothermal amplification | detection or not detection |  | 2,750/sample |  | 12043-60110 | Chemical Metrology Department | 2750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
17 | Nitrogen Dioxide Analyzer | 30 - 3,000 umol/mol | ? 1.5% | 14,500/Item | 3-point calibration | 12050-11701 | Chemical Metrology Department | 14500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
21 | Ozone Analyzer | 0 - 1000 nmol/mol | As specification | 10,500/item | 4-point calibration including zero point | 12050-15201 | Chemical Metrology Department | 10500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
22 | Ozone Transfer Standard | 0 - 1,000 nmol/mol | 2.3% relative | 55,000/Item | EPA: multipoint calibration (min 7-point) | 12050-15202 | Chemical Metrology Department | 55000 | Electrical Metrology Department | 4 | standard, solution, mass, set, ph
303 | Aluminium cylinder - CGA 660 | Cylinder Volume: 10 Liters Strainless Valve: CGA 660 |  | 15,000/cyl | Lot 2563 | 12050-20002 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
198 | Oxygen in Nitrogen | 100 µmol/mol -5 cmol/mol | 1.0 %relative | 25,400/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12050-20109 | Chemical Metrology Department | 25400 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
200 | Carbon Dioxide in Nitrogen | 0.005-20 cmol/mol | 0.5 %relative | 30,000 refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12050-20201 | Chemical Metrology Department | 30000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
219 | Carbon Dioxide in Argon | 100 - 500,000 µmol/mol | 1.0 %relative | 20,000/ cylinder | Cylinder volume: 10 Liters, Filling Pressure :100 bar CGA Valve: 580 | 12050-20205 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
204 | Methane in Air | 0.05-2.5 cmol/mol | 1.0 %relative | 24,100/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12050-20305 | Chemical Metrology Department | 24100 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
205 | Carbon Monoxide in Air | 50 - 65000 µmol/mol | 1.0 %relative | 24,100/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12050-20401 | Chemical Metrology Department | 24100 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
206 | Carbon Monoxide in Nitrogen | 5 - 200 000 µmol/mol | 1.0 - 2.0 %relative | 20,000/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: 100-120 bar CGA Valve: 580 | 12050-20403 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
210 | Propane in Air | 0.5-1 cmol/mol | 1.0 %relative | 24,100/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12050-20605 | Chemical Metrology Department | 24100 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2555 | Nitrogen monoxide in Nitrogen | TRM-E-3071 | 2% relative | 31,000/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 660 | 12050-20702 | Chemical Metrology Department | 31000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
2556 | Nitrous oxide in Nitrogen | TRM-E-3081 | 5% relative | 31,000/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-20801 | Chemical Metrology Department | 31000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
231 | Ethanol in Air | 50 - 200 µmol/mol | 1.0 %relative | 35,000/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-110 bar CGA Valve: 580 | 12050-20901 | Chemical Metrology Department | 35000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
212 | Oxygen, Carbon monoxide, Methane and Hydrogen sulfide in Nitrogen | O2 : 18 cmol/mol CH4: 2.2 cmol/mol CO: 100 µmol/mol H2S: 25 µmol/mol | 1.0 %relative for O2, CH4, CO 5.0 %relative for H2S | 35,000/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: 100-120 bar CGA Valve: 660 | 12050-21000 | Chemical Metrology Department | 35000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
213 | Oxygen, Carbon monoxide and Methane in Nitrogen | O2 : 18 cmol/mol CH4: 2.2-2.5 cmol/mol CO: 100 µmol/mol | 1.0 %relative | 30,000/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12050-21001 | Chemical Metrology Department | 30000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
215 | Iso-butane in Air | 0.1- 1 cmol/mol | 1% relative | 24,500/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12050-21201 | Chemical Metrology Department | 24500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2557 | Carbon dioxide and Methane in Nitrogen | TRM-E-3134 | ?2 %relative ?4%relative | 29,000/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-21301 | Chemical Metrology Department | 29000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
217 | Hydrogen in Nitrogen | 100 - 20 000 µmol/mol | 1.0 %relative | 20,000/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: 100-120 bar CGA Valve: 580 | 12050-21501 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
2553 | Hydrogen in Nitrogen | TRM-E-3150 | ?1 %relative | 20,000/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-21501 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
220 | Hydrogen in Helium | 100 - 20,000 µmol/mol | 1.0% relative | 25,000/ cylinder | Cylinder volume: 10 Liters, Filling Pressure :100 bar CGA Valve: 580 | 12050-21502 | Chemical Metrology Department | 25000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2554 | Hydrogen in Nitrogen | TRM-E-3151 | ?1 %relative | 25,000/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-21502 | Chemical Metrology Department | 25000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
222 | Isobutylene in Air | 100 µmol/mol | 1.0 %relative | 24,500/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: 100-120 bar CGA Valve: 580 | 12050-21801 | Chemical Metrology Department | 24500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
223 | Carbon Dioxide, Oxygen in Nitrogen | 1 - 50 cmol/mol | 1.0 %relative | 28,000/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: Depend on composition CGA Valve: 580 | 12050-21901 | Chemical Metrology Department | 28000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
233 | Carbon dioxide, Nitrous oxide and Oxygen in Nitrogen | CO2: 5 cmol/mol N2O: 30 cmol/mol O2: 40 cmol/mol | 1.0 %relative 1.0 %relative 1.0 %relative | 25,500/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: approx 50 bar CGA Valve: 580 | 12050-22001 | Chemical Metrology Department | 25500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
234 | Isoflurane, Carbon dioxide, Nitrous oxide and Oxygen in Nitrogen | Isoflurane: 0.5 - 2 cmol/mol CO2: 5 cmol/mol N2O: 30 cmol/mol O2: 40 cmol/mol | 1.0% relative 1.0% relative 1.0% relative 1.0% relative | 25,500/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: Depend on composition (6-25 bar), CGA Valve: 580 | 12050-22002 | Chemical Metrology Department | 25500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
235 | Sevoflurane, Carbon dioxide, Nitrous oxide and Oxygen in Nitrogen | Sevoflurane : 0.5-2 cmol/mol CO2 : 5 cmol/mol N2O : 30 cmol/mol O2 : 40 cmol/mol | 1.0% relative 1.0% relative 1.0% relative 1.0% relative | 25,500/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: Depend on composition (6-20 bar), CGA Valve: 580 | 12050-22003 | Chemical Metrology Department | 25500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
236 | Isoflurane in Nitrogen | 0.5 - 2 cmol/mol | 1.0 %relative | 13,400/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: Depend on composition (6-25 bar), CGA Valve: 580 | 12050-22101 | Chemical Metrology Department | 13400 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
237 | Sevoflurane in Nitrogen | Sevoflurane : 0.5-2 cmol/mol | 1.0% relative | 15,000/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: Depend on composition (6-20 bar), CGA Valve: 580 | 12050-22201 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
221 | Methane, Nitrogen in Hydrogen | 500 - 500,000 µmol/mol | 1.0% relative | 27,000/ cylinder | Cylinder volume: 10 Liters, Filling Pressure :100 bar CGA Valve: 580 | 12050-22301 | Chemical Metrology Department | 27000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
224 | Water in Nitrogen | 10 µmol/mol | 10% relative | 36,000/refill per 1 cylinder | Cylinder volume : 30 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-22401 | Chemical Metrology Department | 36000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
2552 | Helium in Argon | TRM-E-3250 | ?1 %relative | 25,400/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-22501 | Chemical Metrology Department | 25400 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2558 | Carbon monoxide Carbon dioxide and Methane in Argon | TRM-E-3251 | 5 %relative | 56,500/refill per 1 cylinder | Cylinder volume : 30 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-22502 | Chemical Metrology Department | 56500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
2551 | Hydrogen, Nitrogen and Methane in Helium | TRM-E-3270 | ?1 %relative | 33,000/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-22701 | Chemical Metrology Department | 33000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
2559 | Benzene / Chloroform / Dichloromethane / Trichloroethylene / Tetrachloroethylene / 1, 2-Dichloroethane / 1, 2-Dichloropropane / 1, 3-Butadiene / Vinyl chloride / Balance : Nitrogen | TRM-E-3280 | ? 5 %relative ? 7 %relative for Vinyl chloride, 1,3-Butadiene, | 140,000/refill per 1 cylinder | Cylinder volume : 10 Liters Filling Pressure : 100 bar CGA Valve: 580 | 12050-22801 | Chemical Metrology Department | 140000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
279 | Oxygen in Nitrogen | 0.01 -5 cmol/mol | 0.7%relative | 30,500/cylinder | Cylinder Volume: approx. 10 Liters | 12050-60102 | Chemical Metrology Department | 30500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
280 | Oxygen in Nitrogen | 10-100 µmol/mol | 1.2%relative | 30,500/cylinder | Cylinder Volume: approx. 10 Liters | 12050-60103 | Chemical Metrology Department | 30500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
292 | Carbon monoxide, Methane, Carbon dioxide in Nitrogen | 10 - 1,000 µmol/mol | ?0.5 %relative | 34,000/cylinder | Cylinder Volume: approx. 10 Liters. In the same concentration | 12050-61301 | Chemical Metrology Department | 34000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
293 | Ammonia in Nitrogen | 20-100 µmol/mol | ?2.5 %relative | 30,600/ cylinder | Cylinder Volume: approx. 10 Liters | 12050-61401 | Chemical Metrology Department | 30600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
294 | Hydrogen in Nitrogen | 100 - 20 000 µmol/mol | ?0.5 %relative | 30,500/cylinder | Cylinder Volume: approx. 10 Liters | 12050-61501 | Chemical Metrology Department | 30500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
295 | Nitrogen Dioxide in Nitrogen | 30 - 3 000 umol/mol | ?2.0 %relative | 27,600/cylinder | Cylinder Volume: approx. 10 Liters | 12050-61701 | Chemical Metrology Department | 27600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
298 | Purity assessment of high purity nitrogen Measurand: O2, CO, CO2, CH4, Ar, H2O | 0.1-100 µmol/mol | 2-30%relative | 16,000/cylinder |  | 12050-66001 | Chemical Metrology Department | 16000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
299 | Purity assessment of high purity oxygen Measurand: N2, CO, CO2, CH4, Ar, H2O | 0.1-100 µmol/mol | 2-30%relative | 18,000/cylinder |  | 12050-66101 | Chemical Metrology Department | 18000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
300 | Purity assessment of high purity carbon monoxide Measurand: O2, N2, CO2, CH4, H2O | 0.1-100 µmol/mol | 2-30%relative | 13,000/cylinder |  | 12050-66201 | Chemical Metrology Department | 13000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
301 | Purity assessment of high purity methane Measurand: O2, CO, CO2, THC, H2O | 0.1-100 µmol/mol | 2-30%relative | 13,000/cylinder |  | 12050-66301 | Chemical Metrology Department | 13000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
14 | Sulfur Dioxide Analyzer | 30 - 5,000 umol/mol | ? 1.0% | 14,500/Item | 3-point calibration | 12051-10501 | Chemical Metrology Department | 14500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
15 | Propane Detector | 0.85 cmol/mol | ? 1.0% | 3,000/item | 1-point calibration | 12051-10601 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
16 | Nitrogen Monoxide Analyzer | 30 - 5,000 umol/mol | ? 1.0% | 14,500/Item | 3-point calibration | 12051-10701 | Chemical Metrology Department | 14500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
18 | Iso-butane Detector | 0.75 cmol/mol | ? 1.0% | 3,000/item | 1-point calibration | 12051-11201 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
19 | Multi Gas Detectors for Carbon Monoxide, Methane, Oxygen and Hydrogen Sulfide | CO: 100 µmol/mol CH4: 2.2 cmol/mol O2 : 18 cmol/mol H2S: 25 µmol/mol | ? 1.0% | 6,000/item | 1-point calibration | 12051-11301 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
20 | Ammonia Detector | 50 µmol/mol | ? 1.0% | 3,000/item | 1-point calibration | 12051-11401 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
302 | Aluminium cylinder 10L- CGA 580 | Cylinder Volume: 10 Liters Brass Valve: CGA 580 |  | 13,000/cyl | Lot 2565 | 12051-20001 | Chemical Metrology Department | 13000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Aluminium cylinder 10L- CGA 660 |  |  | 15,000/cyl |  | 12051-20002 | Chemical Metrology Department | 15000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Aluminium cylinder 30L- CGA 580 |  |  | 20,000/cyl |  | 12051-20003 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Aluminium cylinder 30L- CGA 660 |  |  | 20,000/cyl |  | 12051-20004 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Aluminium cylinder 2.9L- CGA 580 |  |  | 6,000/cyl |  | 12051-20005 | Chemical Metrology Department | 6000 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
196 | Oxygen in Nitrogen | 5-30 cmol/mol | 0.25 %relative | 29,000/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12051-20101 | Chemical Metrology Department | 29000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
197 | Oxygen in Nitrogen | 5-30 cmol/mol | 0.5 %relative | 18,000/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12051-20103 | Chemical Metrology Department | 18000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
232 | Oxygen in Nitrogen | 30 - 100 cmol/mol | 0.5 %relative | 23,000/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: 100-120 bar CGA Valve: 580 | 12051-20104 | Chemical Metrology Department | 23000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
199 | Oxygen in Nitrogen | 10 - 100 µmol/mol | 2.0 %relative | 25,400/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12051-20111 | Chemical Metrology Department | 25400 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
207 | Sulfur Dioxide in Nitrogen | 30 -5 000 µmol/mol | 1.0% relative | 31,000/refill per cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 660 | 12051-20501 | Chemical Metrology Department | 31000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
208 | Propane in Nitrogen | 0.001- 1 cmol/mol | 0.7% relative | 31,000/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12051-20601 | Chemical Metrology Department | 31000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
209 | Propane in Nitrogen | 0.001- 1 cmol/mol | 1.0 %relative | 20,500 refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12051-20603 | Chemical Metrology Department | 20500 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
211 | Nitrogen Monoxide in Nitrogen | 30 - 5 000 µmol/mol | 1.0% relative | 31,000/refill per cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 660 | 12051-20701 | Chemical Metrology Department | 31000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
214 | Hydrogen Sulfide in Nitrogen | 20 - 100 µmol/mol | 4% relative | 27,400/refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 660 | 12051-21101 | Chemical Metrology Department | 27400 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
216 | Ammonia in Nitrogen | 20-100 µmol/mol | 4.0 %relative | 27,400/refill per 1 cylinder | Cylinder Volume: approx. 10 Liters Filling Pressure: 100-120 bar CGA Valve: 660 | 12051-21401 | Chemical Metrology Department | 27400 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
218 | Nitrogen Dioxide in Nitrogen | 30 - 3,000 umol/mol | 2.0% relative | 31,000/refill per cylinder | Cylinder Volumer: approx. 10 Liters Filling Pressure: 100-120 bar CGA Valve: 660 | 12051-21701 | Chemical Metrology Department | 31000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
278 | Oxygen in Nitrogen | 5-30 cmol/mol | 0.25%relative | 23,000/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60101 | Chemical Metrology Department | 23000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
284 | Carbon Monoxide in Nitrogen | 0.005 - 20 cmol/mol | 0.5%relative | 24,000/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60402 | Chemical Metrology Department | 24000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
285 | Sulfur Dioxide in Nitrogen | 100 -5 000 µmol/mol | 0.7%relative | 27,600/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60501 | Chemical Metrology Department | 27600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
286 | Sulfur Dioxide in Nitrogen | 30 - 100 µmol/mol | 1.2% relative | 27,600/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60502 | Chemical Metrology Department | 27600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
287 | Propane in Nitrogen | 0.001- 1 cmol/mol | 0.7%relative | 24,000/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60601 | Chemical Metrology Department | 24000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
288 | Nitrogen Monoxide in Nitrogen | 50 - 5 000 µmol/mol | 0.7%relative | 27,600/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60701 | Chemical Metrology Department | 27600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
2542 | Nitrogen monoxide in Nitrogen | 10 -5,000 ?mol/mol | (0.5-0.7)% relative | 27,600/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60701 | Chemical Metrology Department | 27600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
289 | Nitrous Oxide in Nitrogen | 0.05-10 cmol/mol | 0.7%relative | 27,600/cylinder | Cylinder Volume: approx. 10 Liters | 12051-60801 | Chemical Metrology Department | 27600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
290 | Carbon monoxide, Carbon dioxide in Nitrogen | 1-15 cmol/mol | 0.5%relative | 28,800/cylinder | Cylinder Volume: approx. 10 Liters | 12051-61001 | Chemical Metrology Department | 28800 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
291 | Hydrogen Sulfide in Nitrogen | 25 - 100 µmol/mol | 2.5% relative | 30,600/cylinder | Cylinder Volume: approx. 10 Liters | 12051-61101 | Chemical Metrology Department | 30600 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
296 | Particulate Matter as PM2.5, PM10 (Analysis) | 0 - 500 µg/m3 | As air volume specification (³2.0 %relative) | 2,700/3 filter samples-47mm | By gravimetric analyis of filter sample | 12051-65001 | Chemical Metrology Department | 2700 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
297 | Organic and Elemental Carbon (OC/EC) in Particulate Matter | 0 - 200 µg C/m3 | As air volume specification (³2.0 %relative) | 2,100/3 filter samples | By thermal-optical analyis of filter sample | 12051-65101 | Chemical Metrology Department | 2100 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
7 | Oxygen Detector | 18 - 23 cmol/mol | ? 1.0% | 3,000/item | 1-point calibration | 12052-10102 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
9 | Carbon Dioxide Detector | 0.005-20 cmol/mol | ? 1.0% | 3,000/item | 1-point calibration | 12052-10202 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
11 | Methane Detector | 2.2 -2.5 cmol/mol | ? 1.0% | 3,000/item | 1-point calibration | 12052-10302 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
13 | Carbon Monoxide Detector | 50 - 65000 µmol/mol | ? 1.0% | 3,000/item | 1-point calibration | 12052-10402 | Chemical Metrology Department | 3000 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
201 | Carbon Dioxide in Nitrogen | 0.005-20 cmol/mol | 1.0 %relative | 20,000 refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12052-20203 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
202 | Methane in Nitrogen | 0.005-50 cmol/mol | 0.5%relative | 30,000 refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12052-20301 | Chemical Metrology Department | 30000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
203 | Methane in Nitrogen | 0.005-50 cmol/mol | 1.0%relative | 20,000 refill per 1 cylinder | Cylinder volume: 10 Liters, Filling Pressure :100-120 bar CGA Valve: 580 | 12052-20303 | Chemical Metrology Department | 20000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
283 | Carbon Monoxide in Nitrogen | 5-100 µmol/mol | 1.0%relative | 24,000/cylinder | Cylinder Volume: approx. 10 Liters | 12052-60401 | Chemical Metrology Department | 24000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
6 | Oxygen Analyzer | 0.001 -30 cmol/mol | ? 1.0% | 11,500/item | 3-point calibration | 12053-10101 | Chemical Metrology Department | 11500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
8 | Carbon Dioxide Analyzer | 0.005-20 cmol/mol | ? 1.0% | 11,500/item | 3-point calibration | 12053-10201 | Chemical Metrology Department | 11500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
10 | Methane Analyzer | 0.005-20 cmol/mol | ? 1.0% | 11,500/item | 3-point calibration | 12053-10301 | Chemical Metrology Department | 11500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
12 | Carbon Monoxide Analyzer | 50 - 65000 µmol/mol | ? 1.0% | 11,750/item | 3-point calibration | 12053-10401 | Chemical Metrology Department | 11750 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
282 | Methane in Nitrogen | 0.05-50 cmol/mol | 0.5%relative | 24,000/cylinder | Cylinder Volume: approx. 10 Liters | 12053-60301 | Chemical Metrology Department | 24000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
281 | Carbon Dioxide in Nitrogen | 0.05-20 cmol/mol | 0.5%relative | 24,000/cylinder | Cylinder Volume: approx. 10 Liters | 12054-60201 | Chemical Metrology Department | 24000 | Chemical Metrology Department | 3 | nitrogen, oxygen, carbon, dioxide, monoxide
190 | Chloramphenicol in freeze dried prawn | 0.74 ng/g | 12% relative | 4,000 baht/bottle |  | 1211-20610 | Chemical Metrology Department | 4000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
7 | Luminous flux of a tungsten lamp | 5 lm to 9 000 lm | 0.0089 | 18000 | Gonio-photometer (CCT: 2 000 K to 3 400 K) | 13010-11031 | Thermometry Metrology Department | 18000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
8 | Luminous flux of an LED | 5 lm to 9 000 lm | 0.012 | 18000 | Gonio-photometer (Geometric measurement conditions: full, peak wavelength: white) | 13010-11032 | Thermometry Metrology Department | 18000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
55 | Spectral radiant flux of a tungsten lamp | 0.001 W/nm to 0.07 W/nm | 3.4% to 2.8% (360 nm to 395 nm) 2.8% to 3.6% (395 nm to 830 nm) | 20,000 (+ 500 for total radiant flux) | Gonio-spectroradiometer and integrating sphere (Geometric measurement conditions: full) | 13010-13051 | Thermometry Metrology Department | 20000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
56 | Total radiant flux of a tungsten lamp | 5 W to 20 W | 0.016 | 20000 | Gonio-spectroradiometer and integrating sphere (Geometric measurement conditions: full, wavelength range: 360 nm to 830 nm) | 13010-15080 | Thermometry Metrology Department | 20000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
1 | Luminous intensity of a tungsten lamp | 10 cd ? I? ? 3 000 cd 2 000 K ? CCT ? 3 400 K | 0.98% | 15000 | Detector-based method against a set of reference photometer | 13011-11011 | Thermometry Metrology Department | 15000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
6 | Illuminance responsivity of an illuminance meter or a lux meter | 2 lx ? E? ? 70 lx (range 1) 70 lx ? E? ? 1 400 lx (range 2) 1 200 lx ? E? ? 10 000 lx (range 3) 10 000 lx < E? ? 30 000 lx (range 4) 2 000 K ? CCT ? 3 400 K | 0.91% (range 1, 2, 3) 1.2% (range 4, 5) varied with illuminance level | 5,000 for the first range + 2,500 for any additional range 2, 3, 4 (5 calibration points per range) | Detector-based comparison against reference photometers | 13011-11021 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 7 | meter, power, phase, ac, wattmeter
2513 | Illuminance responsivity of an illuminance meter or a lux meter (LED source) | 1 000 lx < Eu <120 000 lx5 780 K < CCT < 5 930 K | 1.2% (1000 lx ? Eu? 10000 lx)1.6% (10000 lx < Eu ? 120000 lx)varied with illuminance level | 15,000 (15 calibration points) | Detector-based comparison against reference photometers using LED source | 13011-11022 | Thermometry Metrology Department | 15000 | Thermometry Metrology Department | 7 | meter, power, phase, ac, wattmeter
14 | Luminance responsivity of a luminance meter | 0.100 0 cd/m2 ? L? ? 43 000 cd/m2 2 000 K ? CCT ? 3 400 K | 0.0091 | 30,000 for a set of 30 calibration points covering 5 decades | Detector-based comparison against a set of two reference photometers with precision aperture | 13011-11060 | Thermometry Metrology Department | 30000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
13 | Luminance of a luminance source | 0.100 0 cd/m2 ? L? ? 199 900 cd/m2 2 000 K ? CCT ? 6 500 K | 0.011 | 10,000 for the first calibration point + 8,000 for any additional point | Direct measurement by NIMT’s reference luminance meter | 13012-11051 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
15 | Luminance responsivity of a luminance meter | 1 cd/m2 ? L? ? 49 000 cd/m2 3 000 K ? CCT ? 6 000 K | 0.017 | 4,000 for the first calibration point + 1,500 for any additional point within the same range and approximately same CCT (± 50 K) | Comparison method against the reference luminance meter under the solar spectrum uniform source system | 13012-11060 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
2 | Luminous intensity of an incandescent lamp | 10 cd ? I? ? 3 000 cd 2 000 K ? CCT ? 3 400 K | 0.66% | 18,750 for the first lamp + 8,500 for any additional lamp of the same type | Substitution method against two national reference standard lamps | 13015-11011 | Thermometry Metrology Department | 18750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
4 | Illuminance responsivity of a photo-detector head | 12 lx ? E? ? 300 lx 2 000 K ? CCT ? 3 400 K | 0.70% | 6250 | Source-based measurements using national reference standard lamps applying the inverse-square law. | 13015-11021 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
9 | Luminous flux of a tungsten lamp | 5 lm to 9 000 lm | 1.10% | 13,000 for the first lamp + 7,000 for any additional lamp of the same type | Integrating sphere (CCT: 2 000 K to 3 400 K) | 13015-11031 | Thermometry Metrology Department | 13000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
10 | Luminous flux of an LED | 5 lm to 9 000 lm | 1.30% | 13,000 for the first lamp + 7,000 for any additional lamp of the same type | Integrating sphere (Geometric measurement conditions: full, peak wavelength: white) | 13015-11032 | Thermometry Metrology Department | 13000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
12 | Luminous flux of a linear fluorescent lamp | 800 lm to 4 000 lm | 1.20% | 13,000 for the first lamp + 7,000 for any additional lamp of the same type | Integrating sphere (CCT: 2 600 K to 6 200 K) | 13015-11033 | Thermometry Metrology Department | 13000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
17 | Correlated color temperature of a general source | 2 600 K to 3 200 K (tungsten lamp) 2 600 K to 7 500 K (discharge lamp, LED) | 19 K (tungsten lamp) 23 K to 79 K (discharge lamp, LED) | 6500 | Spectroradiometer and integrating sphere | 13015-15021 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
66 | Spectral diffuse/hemispherical reflectance of a spectrally-neutral material | 30 % to 100 % | 2.2 % to 1.4 % (diffuse: 400 nm to 470 nm hemispherical: 400 nm to 430 nm) 1.4 % to 1.3 % (diffuse: 470 nm to 700 nm hemispherical: 430 nm to 700 nm) | 10,000 | Spectrophotometer | 13015-34051 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
3 | Luminous intensity of an incandescent lamp | 10 cd ? I? ? 3 000 cd 2 000 K ? CCT ? 3 400 K | 0.83% | 12,500 for the first lamp + 6,250 for any additional lamp of the same type | Substitution method against two working standard lamps | 13016-11011 | Thermometry Metrology Department | 12500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
5 | Illuminance responsivity of a photo-detector head | 12 lx ? E? ? 300 lx 2 000 K ? CCT ? 3 400 K | 0.80% | 3130 | Source-based measurements using working standard lamps applying the inverse-square law. | 13016-11021 | Thermometry Metrology Department | 3130 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
18 | Regular spectral transmittance of a transmitting filter or film. Measured transmittance values will be reported at 10 nm interval from 200 nm to 900 nm and 50 nm interval from 900 nm to 2 500 nm. | 31.6% < %T ? 97.7% (0.01 ? OD < 0.5) | 200 nm to 2000 nm: Greater than or equal to 0.48% | 6,250/sample for full range (UV-VIS-NIR) 3,750/sample for half range (UV-VIS or VIS-NIR) 3,750/sample | By comparison against a reference ND filters using the reference spectrophotometer. Measurement uncertainties shown are percent relative of transmittance values and they are varied by wavelengths and dependent of UCC. | 13025-14011 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
19 | Diffuse spectral transmittance of a diffusely transmitting material | 380 nm to 780 nm | For 20% haze plate, typical uncertainties in diffuse spectral transmittance are 0.13% to 0.18% throughout the spectrum | 4,000/sample | Direct measurement by single-beam reference spectrophtometer | 13025-14021 | Thermometry Metrology Department | 4000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
20 | Specular excluded spectral diffuse reflectance of a glossy color sample (8o:de) Measured diffused reflectance values will be reported at 5 nm interval from 380 nm to 780 nm. | 0% < %R ? 10% | 0.0003 | 6,250 for the first piece + 3,750 for any additional piece Fee will be scaled down for selected partial wavelength range. | By comparison against a reference glossy ceramic color standard using the reference spectrophotometer with integrating sphere accessory. Measurement uncertainties shown are percent relative of reflectance values and they are varied by wavelengths and dependent of UCC. | 13025-14051 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
24 | Regular spectral reflectance | 5 ? %R ? 100% 200 nm ? l ? 2 500 nm at 5 nm interval | 0.32% to 0.42% for the wavelength range from 200 nm to 885 nm 1.0% to 1.4% for the wavelength range from 890 nm to 2 500 nm *Typical values for a high reflective aluminum mirror. Uncertainties for other materials are likely higher. | 11,250 for full-range. Price for partial-range will be scaled down accordingly. | Absolute method by a double-subtractive double-beam spectrophotometer (Perkin-Elmer Lambda 1050) equipped with a variable-angle reflectance accessory (URA) | 13025-14061 | Thermometry Metrology Department | 11250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
22 | Specular included spectral diffuse reflectance of a glossy color sample (8o:di) Measured diffused reflectance values will be reported at 5 nm interval from 380 nm to 780 nm. | 0% < %R ? 10% | 0.0030 to 0.0022 | 6,250 for the first piece + 3,750 for any additional piece Fee will be scaled down for selected partial wavelength range. | By comparison against a reference glossy ceramic color standard using the reference spectrophotometer with integrating sphere accessory. Measurement uncertainties shown are percent relative of reflectance values and they are varied by wavelengths and dependent of UCC. | 13025-14071 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
25 | Wavelength calibration of a transmitted wavelength standard (holmium oxide solution, holmium oxide filter, dydimium oxide filter etc.) | 200 nm to 2 500 nm | 200 nm to 700 nm: 0.19 nm | 2,500 for first peak identification + 1,250 for any additional peak | Measurement by reference spectrometer | 13025-14150 | Thermometry Metrology Department | 2500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
26 | Wavelength calibration of a reflectance wavelength standard (holmium oxide doped ceramic tile etc.) | 200 nm to 2 500 nm | 200 nm to 700 nm: 0.19 nm | 3,750 for first peak identification + 1,250 for any additional peak | Measurement by reference spectrometer with integrating sphere accessory | 13025-14151 | Thermometry Metrology Department | 3750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
57 | Surface color of general reflecting material, x, y, Y, u', v', L*, a*, b* | - | Varied by parameter of interest and color of UUC, please contact a laboratory staff for full details. | For general material: 7,500 THB for the first piece and 4,000 THB for an additional piece. For a set of color standards: 7,500 THB for the first piece and 2,000 THB for an additional piece. | Spectral diffuse reflectance measurement by single-beam substitution method against a reference color standard | 13025-16010 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
59 | Transmitted color of a transmitting material | 360 nm to 860 nm | Values and uncertainties vary depending on sample's color. | 3,500/sample | Direct measurement by single-beam reference spectrophtometer | 13025-16030 | Thermometry Metrology Department | 3500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
60 | Haze and luminous transmittance of a diffusely transmitting material | 380 nm to 780 nm | For 20% haze plate, typical uncertainty in transmission haze is 0.16%. | 5,000/sample | Based on ASTM D1003 standard method | 13025-16070 | Thermometry Metrology Department | 5000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
61 | Whiteness and Tint | 380 nm to 780 nm W greater than 40 and less than 5Y-280 | 1.5 | 6,500/sample | Spectral diffuse reflectance measurement by single-beam substitution method against a reference white diffuse reflectance standard | 13025-16120 | Thermometry Metrology Department | 6500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
74 | ND Filter (OD0.1, OD0.4, OD1.0) | 380 nm to 780 nm | Spectral regular transmittance 0.74% to 0.86% for OD0.1 0.86% to 1.2% for OD0.4 T<10%, Uabs = 0.29% and T³10%, Urel = 3.1% for OD1.0 | 10,000/sample | Comparison using the reference spectrophotometer. | 13025-34011 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
75 | Holmium oxide glass filter | 240 nm to 640 nm | Uabs = 0.20 nm | 17,500 for 4 peaks identification + 1,250 for any additional peak | Comparison using the reference spectrophotometer. | 13025-34150 | Thermometry Metrology Department | 17500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
27 | Regular transmittance responsivity of transmission meter and tint meter | 0% < %T ? 100% | 2% | 6250 | By comparison against working standard ND filters at the nominal optical density values of 0.1, 0.3, 0.5 and 1.0. | 13026-14011 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 7 | meter, power, phase, ac, wattmeter
21 | Specular excluded spectral diffuse reflectance of a matte color sample (8o:de) Measured diffused reflectance values will be reported at 5 nm interval from 380 nm to 780 nm. | 0% < %R ? 10% | 0.0003 | 6,250 for the first piece + 3,750 for any additional piece Fee will be scaled down for selected partial wavelength range. | By comparison against a reference matte ceramic color standard using the reference spectrophotometer with integrating sphere accessory. Measurement uncertainties shown are percent relative of reflectance values and they are varied by wavelengths and dependent of UCC. | 13026-14051 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
23 | Specular included spectral diffuse reflectance of a matte color sample (8o:di) Measured diffused reflectance values will be reported at 5 nm interval from 380 nm to 780 nm. | 0% < %R ? 10% | 0.0030 to 0.0022 | 6,250 for the first piece + 3,750 for any additional piece Fee will be scaled down for selected partial wavelength range. | By comparison against a reference matte ceramic color standard using the reference spectrophotometer with integrating sphere accessory. Measurement uncertainties shown are percent relative of reflectance values and they are varied by wavelengths and dependent of UCC. | 13026-14071 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
58 | Surface Color Responsivity of Colorimeter, x, y, Y, u', v', L*, a*, b* | 380 nm to 780 nm | Varied by parameter of interest and color of standard. | 6250 | Direct measurement by using the basic color checked standard | 13026-16010 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
31 | Power responsivity of a trap detector | Power level: < 1 mW | 0.026% (power responsivity) 0.087% (spectral responsivity) | 35,000 per wavelength point + 10,000 for spectral responsivity | - The power responsivity is calibrated against a cryogenic radiometer at laser wavelengths (406.7 nm, 413.1 nm, 476.2 nm, 530.9 nm, 568.2 nm, 632.8 nm, 647.1 nm, 799.3 nm) - Spectral responsivity scale is interpolated between the laser wavelengths based | 13030-12040 | Thermometry Metrology Department | 35000 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
2512 | Spectral responsivity of a broadband detector | Wavelength: 400 nm ? l ? 800 nmPower level: 4.8 nW to 2.6 mW | 0.25% to 0.22% | 20,000 for 10 nm step, 30,000 for 5 nm step | Comparison agianst two reference trap-detectors under monochromatic radiation | 13031-12011 | Thermometry Metrology Department | 20000 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
46 | Measurement of linearity response of a photo-detector | At discrete wavelengths from 200 nm to 2 500 nm Beam size ranges from 25 micron to 5 mm | 200 nm to 400 nm: ± 0.1% 400 nm to 2 500 nm: ± 0.05% | 4,375 for one combination of wavelength and beam size of one photo-detector | Flux superposition method | 13032-12900 | Thermometry Metrology Department | 4375 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
47 | Measurement of uniformity of response of a photo-detector | At discrete wavelengths from 200 nm to 2 500 nm Beam size ranges from 25 micron to 5 mm | 200 nm to 400 nm: ± 0.2% 400 nm to 2 500 nm: ± 0.1% | 6,250 for one combination of wavelength and beam size of one photo-detector | x-y linear scan | 13032-12910 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
34 | Intregrated irradiance responsivity of broadband radiometer | 315 nm to 400 nm (UVA ) | 4.40% | 6,250 per probe for the first irradiance level point + 3,130 for any additional irradiance level point | Source based calibration against a reference spectroradiometer using filtered 500 W Hg(Xe) as light source | 13034-12090 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
35 | Intregrated irradiance responsivity of broadband radiometer | 280 nm to 315 nm (UVB) | 6.30% | 6,250 per probe for the first irradiance level point + 3,130 for any additional irradiance level point | Source based calibration against a reference spectroradiometer using filtered 500 W Hg(Xe) as light source | 13034-12091 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
36 | Intregrated irradiance responsivity of broadband radiometer | 280 nm to 400 nm (UVA+UVB) | 6.00% | 6,250 per probe for the first irradiance level point + 3,130 for any additional irradiance level point | Source based calibration against a reference spectroradiometer using filtered 500 W Hg(Xe) as light source | 13034-12092 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
41 | UVC irradiance responsivity of an UVC radiometers | Wavelength: 240 nm to 270 nm Irradiance level: 2 mW?cm-2 to 20 W?m-2 (Standard points: 2, 40, 160, 225, 350, 450, and 500 mW?cm-2) | 3.1%, varied by DUT and irradiance level | 8,750 THB / fisrt point + 2,500 THB / for additional standard point(s) + 4,375 THB / for additional point(s) | Source based calibration against a reference spectroradiometer using low pressure Hg lamps as calibration source with ND filters | 13034-12093 | Thermometry Metrology Department | 8750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
42 | Erythemally Weighted Effective Global Solar UV Irradiance Responsivity of a UVE Radiometer/UV Index Meter | 0.012 W?m-2 ? Eer ? 0.500 W?m-2 250 nm ? l ? 400 nm | 8.8 % to 9.3 % varied with effective irradiance level and DUT | 6,250 for the first point + 3,750 for any additional point | Source based comparison against working standard spectroradiometer by using a properly filtered Xe discharge lamp as the calibration light source | 13034-12095 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Integrated Irradiance Responsivity of a Bilirubin Phototherapy Radiometer | "1.0 mW?cm-2?nm-2 ? Eb (l) ? 20.0 mW?cm-2?nm-2 425 nm ? l ? 475 nm (Based on AAP CPG)" | "2.2 % varied with mean spectral irradiance level and DUT" | 6,250 for the first level point + 3,750 for any additional level poin | "Source based calibration against a working standard spectroradiometer by using a 445 nm blue LED panel light source as the calibration light source" | 13034-12096 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Integrated Irradiance Responsivity of a Bilirubin Phototherapy Radiometer | "50.0 mW?cm-2 ? Eb ? 750.0 mW?cm-2 400 nm ? l ? 500 nm (Other CPGs)" | "2.1 % varied with irradiance level and DUT" | 6,250 for the first level point + 3,750 for any additional level poin | "Source based calibration against a working standard spectroradiometer by using a 445 nm blue LED panel light source as the calibration light source" | 13034-12096 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Integrated Irradiance Responsivity of a Bilirubin Phototherapy Radiometer | "50.0 mW?cm-2 ? Eb ? 1 000.0 mW?cm-2 400 nm ? l ? 550 nm (Based on IEC 60601-2-50)" | "2.2 % varied with irradiance level and DUT" | 6,250 for the first level point + 3,750 for any additional level poin | "Source based calibration against a working standard spectroradiometer by using a 445 nm blue LED panel light source as the calibration light source" | 13034-12096 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
37 | Intregrated irradiance responsivity of broadband radiometer | Arbitary range within 280 nm to 1 020 nm | Varied by selected range | 6,250 per probe for the first irradiance level point + 3,130 for any additional irradiance level point | Source based calibration against a reference spectroradiometer using filtered 500 W Hg(Xe) as light source | 13034-12099 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
43 | Integrated Radiant Exposure Responsivity of a UV Radiometer | 315 nm ? ? ? 400 nm 2 W?m-2 ? Ee ? 40 W?m-2 30 s ? t ? 120 s | 5.0 % to 6.3 % varied with exposure time, irradiance level and DUT | 8,500 for the first point + 4,200 for any additional point | Comparison agiants a working standard spectroradiometer with a calibrated stopwatch applying the filtered 500W Hg(Xe) arc lamp | 13034-12100 | Thermometry Metrology Department | 8500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
44 | Integrated Radiant Exposure Responsivity of a UV Radiometer | 280 nm ? ? ? 400 nm 5 W?m-2 ? Ee ? 40 W?m-2 30 s ? t ? 180 s | 5.5 % to 6.8 % varied with exposure time, irradiance level and DUT | 8,125 for the first point + 3,750 for any additional point | Comparison agiants a working standard spectroradiometer with a calibrated stopwatch applying the 300W Ultra-Vitalux® UV lamp | 13034-12101 | Thermometry Metrology Department | 8125 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
38 | Intregrated irradiance responsivity of narrowband radiometer | 365 nm ± 10 nm | 0.044 | 6,250 per probe for the first irradiance level point + 3,130 for any additional irradiance level point | Source based calibration against a reference spectroradiometer using filtered 500 W Hg(Xe) as light source | 13034-12110 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
39 | Intregrated irradiance responsivity of narrowband radiometer | 254 nm ± 10 nm | 10% |  |  | 13034-12113 | Thermometry Metrology Department | 7880 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
40 | Intregrated irradiance responsivity of narrowband radiometer | Arbitary range within 280 nm to 1 020 nm | Varied by selected range |  |  | 13034-12119 | Thermometry Metrology Department | 7880 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
32 | Solar irradiance responsivity of a Si-pyranometer | Wavelength: 400 nm to 1 100 nm Power Level: 50 W.m-2 to 380 W.m-2 | 0.02 | 6250 | Comparison against a working standard spectroradiometer applying 300 W Xe arc lamp with 1.5 global airmassfilter | 13034-12716 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
33 | Solar irradiance responsivity of a thermopile-detector pyranometer | 280 nm to 2 500 nm (160 W.m2 to 520 W.m2) | 2.1% to 3.4% | 9000 | Soruce based calibration against a working standard spectroradiometer applying a solar simulator system | 13034-12717 | Thermometry Metrology Department | 9000 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
50 | Measurement of absolute / relative spectral irradance of a light source | 250 nm to 1 020 nm | Greater than or equal to 2.4% | 6,250 for the first lamp + 3,130 for any additional lamp of the same type | Measurement by a working standard diode array spectroradiometer | 13034-13013 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
67 | Measurement of absolute / relative spectral irradance of a light source | 250 nm to 1020 nm | 3.4% to 9.8% | 6,250 | Measurement by a working standard diode array spectroradiometer | 13034-33013 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Solar simulator performance testing according to IEC 60904-9 | Data sampling time of DUT ? 100 ms for instability evaluation | 0.05% | 7,500 for LTI and STI 3,500 for LTI only |  | 13034-33014 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Solar simulator performance testing according to IEC 60904-9 | 80% of active area for non-uniformity evaluation | 0.01% | 12,000 for 98 test positions + 500/additional test position |  | 13034-33014 | Thermometry Metrology Department | 12000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
71 | Solar simulator performance testing according to IEC 60904-9 | 400 nm to 1,100 nm (300 nm to 1 200 nm) for spectral match evaluation | 0.001 | For 400 nm to 1 100 nm, 7,500/irradiance level For 300 nm to 1 200 nm, 9,000/irradiance level | Measurements based on IEC 60904-9 Edition 3.0 2020-09 | 13034-33014 | Thermometry Metrology Department | 400 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
51 | Measurement of integrated absolute irradance of a light source | 320 nm to 400 nm (UVA) | 0.042 | 6,250 for the first lamp (per range) + 3,130 for any additional lamp of the same type (per range) | Measurement by a working standard radiometer | 13035-15090 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
52 | Measurement of integrated absolute irradance of a light source | 260 nm to 320 nm (UVB) | 0.064 | 6,250 for the first lamp (per range) + 3,130 for any additional lamp of the same type (per range) | Measurement by a working standard radiometer | 13035-15091 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
53 | Measurement of integrated absolute irradance of a light source | 400 nm to 500 nm | 0.0357 | 6,250 for the first lamp (per range) + 3,130 for any additional lamp of the same type (per range) | Measurement by a working standard radiometer | 13035-15098 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
72 | UVC Intregrated irradiance of customer's source | 220 nm to 280 nm | 0.054 | 6,400 per probe for the first irradiance level point + 2,900 for any additional irradiance level point | Measurement by reference UVC meter | 13035-32093 | Thermometry Metrology Department | 6400 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
68 | Measurement of integrated absolute irradance of a light source | 320 nm to 400 nm (UVA) | 0.042 | 6,130 (per range) | Measurement by a working standard radiometer | 13035-35090 | Thermometry Metrology Department | 6130 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
69 | Measurement of integrated absolute irradance of a light source | 260 nm to 320 nm (UVB) | 0.064 |  |  | 13035-35091 | Thermometry Metrology Department | 7880 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
70 | Measurement of integrated absolute irradance of a light source | 400 nm to 500 nm | 0.0357 |  |  | 13035-35098 | Thermometry Metrology Department | 7880 | Thermometry Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Spectral irradiance of a tungsten- 1 halogen lamp (FEL and Ushio type) | 250 nm ??? 430 nm 430 nm <?? 1 100 nm 100 nm <?? 1 800 nm 1 800 nm <?? 2 400 nm | 3.6% to 2.2% 2.2% to 1.7% 2.1% to 3.2% 3.2% to 6.8% | 11,250 for the first range of the first lamp + 5,630 for the first range of any additional lamp of the same type + 3,750 for any additional range | By scanning measurements against a set of two spectral irradiance standard lamps using a double subtractive monochromator | 13037-13011 | Thermometry Metrology Department | 11250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
49 | Spectral irradiance of a duterium lamp | 200 nm to 400 nm | 5.3% to 12% | 13,750 for the first lamp + 5,630 for any additional lamp | By scanning measurements against a spectral irradiance reference national standard lamp using a double subtractive monochromator | 13037-13012 | Thermometry Metrology Department | 13750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
54 | Spectral radiance of radiance source | Wavelength range: 320 nm to 1 100 nm Bandwidth: 4.0 nm Correlated color temperature (CCT) 2 000 K ? CCT ? 6 500 K | 2.6% to 7.0% | 20,000 for the first calibration point + 5,000 for any additional point | Comparison method for relative spectral measurement and direct measurement for absolute luminance by a reference luminance meter. | 13037-13021 | Thermometry Metrology Department | 20000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
28 | Absolute spectral irradiance responsivity of a spectroradiometer with irradiance probe | 200 nm to 1 700 nm | Greater than or equal to 2.6% | 11,250 for scanning type and 6,250 for diode-array type | Direct measurement against a set of two FEL working standard lamps | 13038-12023 | Thermometry Metrology Department | 11250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
48 | Spectral irradiance responsivity of low-sense UV and blue-wavelength spectroradiometer | Wavelength: 200 nm ? ? ? 500 nm Power level: 0.002 W?m-2?nm-1 to 1 W?m-2?nm-1 | ³ 4.7%, varied by DUT, wavelength, bandwdith and power level | 10,000 / Irradiance level | Comparison against reference spectroradiometer using 300 W Xe lamp | 13038-12024 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
30 | Absolute spectral radiance responsivity | Radiance: 0.000 5 W.sr-1. m-2.nm-1 Type of source: Tungsten-based integrating sphere source | 2.8% to 7.6% | 7,500 for diode array type spectrometer 10,000 for scanning type spectrometer | Direct measurement on reference radiance gauge | 13038-12030 | Thermometry Metrology Department | 7500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
45 | Responsivity of PAR radiometer | 400 nm to 700 nm | 2.00% | 6,250 per probe for the first irradiance level point + 3,130 for any additional irradiance level point | Direct measurement against a set of two FEL working standard lamps | 13038-12120 | Thermometry Metrology Department | 6250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
29 | Relative spectral radiance responsivity of a spectroradiometer with radiance probe | 250 nm to 2 500 nm | Greater than or equal to 2.5% | 11,250 for scanning type and 6,250 for diode-array type | Direct measurement against a set of two FEL working standard lamps | 13039-12030 | Thermometry Metrology Department | 11250 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
62 | Correlated color temperature of a general source | 2 600 K to 3 200 K (tungsten lamp) 2 600 K to 7 500 K (discharge lamp, LED) | 13 K (tungsten lamp) 14 K to 33 K (discharge lamp, LED) | 6,500 (+ 500 for chromaticity coordinate) | Spectroradiometer and photometric bench | 13041-15021 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
64 | Correlated color temperature response of a color temperature meter | 2 600 K to 3 200 K (tungsten source) 2 600 K to 10 000 K (spectrally tunable LED source) | 13 K (tungsten source) 25 K to 133 K (spectrally tunable LED source) | 6,500 for the first calibration point + 1,000 for any additional point | Spectroradiometer | 13041-15030 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
63 | Emitted color of a general source | - | 0.001 (tungsten lamp) 0.002 (discharge lamp, LED) | 6,500 (+ 500 for correlated color temperature) | Spectroradiometer and photometric bench | 13041-15040 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
65 | Chromaticity response of a colorimeter | - | 0.001 (tungsten source) 0.002 (spectrally tunable LED source) | 6,500 for the first calibration point + 1,000 for any additional point | Spectroradiometer | 13041-15050 | Thermometry Metrology Department | 6500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
16 | Averaged luminous intensity of a single-packaged LED | 0.1 cd to 50 cd | 2.4% to 3.3% varied with LED color | 15,000 for the first LED (CIE condition A and B) + 7,500 for any additional LED of the same type | Photometric bench, spectroradiometer and reference lamps | 13045-11012 | Thermometry Metrology Department | 15000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
11 | Luminous flux of a single-packaged LED | 0.1 lm to 100 lm | 2.1% to 2.6% varied with LED color | 12,500 for the first LED + 6,250 for any additional LED of the same type | Integrating sphere (Geometric measurement conditions: full) | 13045-11032 | Thermometry Metrology Department | 12500 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Spectral responsivity of a PV device | Wavelength: 300 nm ? ? ? 1200 nm Power level: 0.2 mW to 4 mW | 2.0%, varied by DUT, wavelength and power level | 10,000 for the first bias light level + 5,000 for additional bias light level | IEC 60904-8 Clause 7 | 13050-12001 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
2511 | Spectral responsivity of a PV device | Wavelength: 300 nm ? ? ? 1200 nm Power level: 0.2 mW to 4 mW | 2.0%, varied by DUT, wavelength and power level | 10,000 for the first bias light level + 5,000 for additional bias light level | IEC 60904-8 Clause 7 | 13050-12001 | Thermometry Metrology Department | 10000 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
73 | Sensitivity of a pyranometer | (600 to 1000) W/m2 | 0.03 | 10000 | Comparison to a reference pyranometer in accordance with ISO 9847:2023 Type B1 | 13050-31001 | Thermometry Metrology Department | 10000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Adjustment Charge |  |  | 25 % of the calibration cost of each range |  | 14000 - 10001 | Acoustic and Vibration | 25 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Measurement Microphone type LS1P by the reciprocity technique (Pressure sensitivity level) | 20 Hz to 10 kHz | 0.04 dB to 0.07 dB | 10,630/piece |  | 14010 - 11101 | Acoustic and Vibration | 10630 | Acoustic and Vibration | 6 | measurement, detector, additional, frequency, irradance
| Measurement Microphone type LS2P by the reciprocity technique (Pressure sensitivity level) | 20 Hz to 20 kHz | 0.04 dB to 0.10 dB | 10,630/piece |  | 14010 - 11102 | Acoustic and Vibration | 10630 | Acoustic and Vibration | 6 | measurement, detector, additional, frequency, irradance
| Measurement Microphone by the insert voltage technique (pressure sensitivity level) | 250 Hz and 1 kHz | 0.08 dB | 5,130/piece |  | 14010 - 11103 | Acoustic and Vibration | 5130 | Acoustic and Vibration | 6 | measurement, detector, additional, frequency, irradance
| Measurement Microphone by the comparison method (free-field sensitivity level) | 31.5 Hz to 16 kHz 74 dB | 0.1 dB to 0.4 dB | 7,880/piece |  | 14010 - 11201 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 6 | measurement, detector, additional, frequency, irradance
| Measurement Microphone type WS2 by the insert voltage technique (microphone sensitivity) and by The electrostatic actuator (frequency respone) | 20 Hz to 20 kHz | Sensitivity : 0.10 dB Frequency Response : 0.10 to 0.30 dB | 7,500/piece |  | 14010 - 11202 | Acoustic and Vibration | 7500 | Acoustic and Vibration | 6 | measurement, detector, additional, frequency, irradance
| Measurement Microphone type WS3 by the insert voltage technique (microphone sensitivity) by The electrostatic actuator or coupler comparison (frequency respone) | 20 Hz to 20 kHz | Sensitivity : 0.10 dB Frequency Response : 0.10 to 0.30 d | 7,500/piece |  | 14010 - 11203 | Acoustic and Vibration | 7500 | Acoustic and Vibration | 6 | measurement, detector, additional, frequency, irradance
| Pistonphone or Sound Calibrator | 74 dB to 124 dB 250 Hz, 1kHz | 0.08 dB | 4,880 first level 1,000/next level |  | 14010 - 12101 | Acoustic and Vibration | 4880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Multifunction Acoustic Calibrator or Multi-frequency Sound Calibrator | 74 dB to 124 dB 31.5 Hz to 16 kHz | 0.09 dB | 16,880/piece |  | 14010 - 12201 | Acoustic and Vibration | 16880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Sound Level Meter | Based on IEC 61672 Part 3 | Class 1, 2 | 10,880/piece |  | 14010 - 13101 | Acoustic and Vibration | 10880 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Sound Level Meter | Based on IEC 61672 Part 3 | Class 1, 2 | 10,880/piece |  | 14010 - 13101 | Acoustic and Vibration | 10880 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Sound Level Meter (Free-field sound pressure response level and absolute sensitivity) | 31.5 Hz to 16 kHz 74 dB | 0.2 dB to 0.4 dB | 7,880/piece |  | 14010 - 13201 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Sound Exposure Meter or Noise Dose Meter for a frequency | range of 63 Hz – 8 kHz |  | 17,500 Baht |  | 14010 - 13301 | Acoustic and Vibration | 17500 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Audiometer with supra-aural earphone for a frequency | 125 Hz - 8 kHz |  | 7,500 Baht |  | 14010 - 14101 | Acoustic and Vibration | 7500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Bone-conduction pure-tone audiometer | 250 Hz - 4 kHz | Type 1, 2, 3 áÅÐ 4 | 7,500/piece |  | 14010 - 14501 | Acoustic and Vibration | 7500 | Chemical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Sound Absorption Coeficients of Materials for low frequency | 50 Hz - 1.6 kHz | 7.10% | 5,000/piece |  | 14010 - 67001 | Acoustic and Vibration | 5000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Sound Absorption Coeficients of Materials for high frequency | 500 Hz - 6.4 kHz | 7.10% | 5,000/piece |  | 14010 - 67002 | Acoustic and Vibration | 5000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Sound Transmission Loss of Materials for frequency | 50 Hz - 6.4 kHz | 3.80% | 8,750 Baht |  | 14010 - 67003 | Acoustic and Vibration | 8750 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Tuning Fork | 500 Hz - 5 kHz | 0.8 Hz | 3,125/piece |  | 14010 - 67101 | Acoustic and Vibration | 3125 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Absorption Coefficient of Material | 80 Hz to 6300 Hz | 1.5% to 4.0% | 3,500/piece |  | 14010 - 99001 | Acoustic and Vibration | 3500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Sound Calibration/Pistonphone | 74 dB to 124 dB, 250 Hz, 1 kHz |  |  |  | 14010-12101 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Sound Level Meter Based on IEC61672 Part 3 | Based on IEC 61672 Part 3 |  |  |  | 14010-13101 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Ear Simulator | 100 Hz - 16 kHz | 100 Hz to 12.5kHz : 0.4 dB >12.5 kHz to 16 kHz : 0.5 dB | 8,750/piece |  | 14010-14401 | Acoustic and Vibration | 8750 | Thermometry Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Dental Scaler (Noise Measurement) Based On ISO 22374 | 20 Hz - 40 kHz | as specification | 3,750/piece |  | 14010-67301 | Acoustic and Vibration | 3750 | Chemical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Vibration Meter | 10 m/s2 to 100 m/s2 50 Hz to 5 kHz | 1.00% | 10,000/piece |  | 14020 - 11101 | Acoustic and Vibration | 10000 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Vibration meter at Low frequency | 1 m/s2 to 100 m/s2 1 Hz to 160 Hz | 1.00% | 10,000/piece |  | 14020 - 11102 | Acoustic and Vibration | 10000 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Vibration Meter for Low frequency by comparison with standard accelerometer. | 1 m/s² to 100 m/s² 10 Hz to 5kH | 1.00% | 15,000/piece |  | 14020 - 11103 | Acoustic and Vibration | 15000 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Vibration Calibrator | 10 m/s2 to 100 m/s2 50 Hz to 5 kHz | 1.00% | 5,000 first point 1,250/next point |  | 14020 - 12101 | Acoustic and Vibration | 5000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Mobile vibration calibrator based on ISO16063-44 | 1 m/s2 to 100 m/s2 40 Hz to 5 kHz | 1.2 % for Acceleration 1 % for Frequency 1 % for Distortion | 35,000 /first 8 points 5,000/next point |  | 14020 - 12102 | Acoustic and Vibration | 35000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by Laser Interferometry (charge sensitivity) | 10 m/s2 to 200 m/s2 40 Hz to 800 Hz 10 m/s2 to 370 m/s2 > 800 Hz to 5 kHz | 0.5 % 0.8 % | 13,750/piece |  | 14020 - 13101 | Acoustic and Vibration | 13750 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by comparison to reference transducer (charge sensitivity) | 10 m/s2 to 100 m/s2 50 Hz to 5 kHz | 1.00% | 8,750 first point 1,250/next point |  | 14020 - 13102 | Acoustic and Vibration | 8750 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Shock Transducer by Laser interferometry | 20 g to 500 g | 1.00% | 11,250 /piece |  | 14020 - 13201 | Acoustic and Vibration | 11250 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Shock Transducer by comparison to reference transducer (shock sensitivity) | 20 g to 10,000 g 0.1 ms - 3 ms | 4.50% | 7,000 first point 1,000/next point |  | 14020 - 13202 | Acoustic and Vibration | 7000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by Laser Interferometry (voltage sensitivity) | 10 m/s2 to 200 m/s2 50 Hz to 800 Hz 10 m/s2 to 370 m/s2 > 800 Hz to 5 kHz | 0.5 % 0.8 % | 13,750/piece |  | 14020 - 13301 | Acoustic and Vibration | 13750 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by Laser Interferometry (Voltage sensitivity in modulus) | 0.3 m/s2 to 10 m/s2 1 Hz to 160 Hz | 1% | 13,000/piece |  | 14020 - 13302 | Acoustic and Vibration | 13000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by Laser Interferometry (Phase shift of Voltage transducer type) | 0.3 m/s2 to 10 m/s2 1 Hz to 160 Hz | 1.0 degree | 13,000/piece |  | 14020 - 13303 | Acoustic and Vibration | 13000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by Laser Interferometry (Voltage sensitivity in modulus with Phase shift) | 0.3 m/s2 to 10 m/s2 1 Hz to 160 Hz | 1.0% and 1.0 degree | 15,000/piece |  | 14020 - 13304 | Acoustic and Vibration | 15000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by comparison to reference transducer (voltage sensitivity) | 0.3 m/s2 to 10 m/s2 1 Hz-80Hz | 1.00% | 8,750/piece |  | 14020 - 13305 | Acoustic and Vibration | 8750 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Accelerometer by comparison to reference transducer (Voltage sensitivity) | 10 m/s2 to 100 m/s2 40 Hz to 5 kHz | 1.00% | 10,000/piece |  | 14020 - 13306 | Acoustic and Vibration | 10000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Laser Vibrometer by comparison to reference vibrometer (modulus) | 10 m/s2 to 50 m/s2 40 Hz to 5 kHz | 0.50% | 50,000/piece |  | 14020 - 13401 | Acoustic and Vibration | 50000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Tachometer (Beam of infrared LED technique) | 900 RPM to 99,000 RPM | 1.00% | 5,000/piece |  | 14020 - 18102 | Acoustic and Vibration | 5000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Stroboscope | 210 RPM to7500 RPM | 1.00% | 5,000/piece |  | 14020 - 18103 | Acoustic and Vibration | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Charge Amplifier | 20 Hz to 10 kHz | 0.10% | 5,380/piece |  | 14020 - 19101 | Acoustic and Vibration | 5380 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Conditioning Amplifier (Voltage Amplifier) | 40 Hz to 5 kHz | 0.10% | 6,250/piece |  | 14020 - 19102 | Acoustic and Vibration | 6250 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Conditioing Amplifier (ICP Amplifier) | 40 Hz to 5 kHz | 0.10% | 6,250/piece |  | 14020 - 19103 | Acoustic and Vibration | 6250 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Electrical Controller | Analog Input Voltage:23 mV to 6 V analog input charge 15 to 7000 pC Frequency: 10 hz to 100 kHz | ?0.5 mV/V for Frequency 10 – 40 Hz ?0.4 mV/V for Frequency 40 Hz – 1 kHz ?0.5 mV/V for Frequency 1 – 20 kHz ?0.8 mV/V for Frequency 20 – 50 kHz ?1.6 mV/V for Frequency 50 – 100 kHz ?0.7 fC/pC for Frequency 10 – 40 Hz ?0.6 fC/pC for Frequency 40 Hz – 1 kHz ?0.7 fC/pC for Frequency 1 – 20 kHz ?1.0 fC/pC for Frequency 20 – 50 kHz ?1.8 fC/pC for Frequency 50 – 100 kHz | 60,000 /piece |  | 14020 - 19104 | Acoustic and Vibration | 60000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Low-frequency vibration sensor | Vertical direction:0.1 Hz to 100 Hz Horizontal direction: 0.1 to 80 Hz | 1.00% | 20,000 /piece |  | 14020-11104 | Acoustic and Vibration | 20000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
2535 | Seismic sensor jointed with digitizer by laser interferometer | Horizontal and Vertical directions: 0.1 to 10 Hz Horizontal and Vertical directions: >10 to 80 Hz | 2.0% 7.0% | 12,500 / piece |  | 14020-11105 | Acoustic and Vibration | 12500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vibration Table áºº onsite | 10 Hz-100 Hz |  |  |  | 14020-12102 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
2533 | Shock accelerometer by comparison to a reference accelerometer based on ISO 16063-22 (Charge sensitivity in modulus) | 2000 m/s2 to 100000 m/s2 Duration Time 0.08 ms to 1.6 ms | 0.03 | 15,000 / first 3 points 3,000/ next point |  | 14020-13202 | Acoustic and Vibration | 15000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
2531 | Shock Accelerometer by laser interferometry (Charge sensitivity in modulus) | 50 m/s2 to 2000 m/s2 Duration Time 1.6 ms to 6.0 ms | 0.01 | 20,000 / first 3 points 4,000/ next point |  | 14020-13203 | Acoustic and Vibration | 20000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Vibration sensor-voltage sensitivity | 40 Hz - 5 kHz |  |  |  | 14020-13306 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Laser Vibrometer base on ISO 16063-41 for low frequency | 0.1 Hz to less than 0.4 Hz 0.4 Hz to 40 Hz more than 40 Hz to 160 Hz | 1.0 % 0.5% 2.0% | 50,000/piece |  | 14020-13402 | Acoustic and Vibration | 50000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Laser Vibrometer base on ISO 16063-41 for medium frequency | 50 Hz to 15 kHz more than 15 kHz to 20 kHz | 0.5% 1.0% | 50,000/piece |  | 14020-13403 | Acoustic and Vibration | 50000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
2534 | Shock accelerometer by comparison to a reference accelerometer based on ISO 16063-22 (Voltage sensitivity in modulus) | 2000 m/s2 to 100000 m/s2 Duration Time 0.08 ms to 1.6 ms | 0.03 | 15,000 / first 3 points 3,000/ next point |  | 14020-13404 | Acoustic and Vibration | 15000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
2532 | Shock accelerometer by laser interferometry (Voltage sensitivity in modulus) | 50 m/s2 to 2000 m/s2 Duration Time 1.6 ms to 6.0 ms | 0.01 | 20,000 / first 3 points 4,000/ next point |  | 14020-13405 | Acoustic and Vibration | 20000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Pendulum Shock (on-site) | 500 m/s2 to 5000 m/s2 | 5% | 31,250/piece |  | 14020-13601 | Acoustic and Vibration | 31250 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Laser Vibrometer | 20 kHz | 3.50% | 4,375/point |  | 14020-18101 | Acoustic and Vibration | 4375 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Displacement (Unloaded scaler), Frequency (Dental scaler) | 0.1 µm to 220.0 µm. 18 kHz to 20 kHz | 5.0% 2.0% | 11,250/piece |  | 14020-18104 | Acoustic and Vibration | 11250 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Tachometer (Engine Ignition Signal with Sparking Plug (Simulation Set)) | 120 RPM to 36,000 RPM | 1% | 5,000 /piece |  | 14020-18105 | Acoustic and Vibration | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2528 | Tachometer and RPM Meter (Vibration Technique) | 120 RPM to 21,000 RPM | 0.01 | 5,000/piece |  | 14020-18106 | Acoustic and Vibration | 5000 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
2529 | GPS/GNSS Speed Calibration Dynamics Test (Field Test) onsite | 10 km/h to 160 km/h | ? 0.2 km/h for speed 10 km/h to 100 km/h ? 0.25 km/h for speed >100 km/h to 160 km/h | 55,000/piece |  | 14020-18107 | Acoustic and Vibration | 55000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Accelerometer | 2 Hz to 50 Hz |  |  |  | 14020-19102 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Safety Belt | 11 Hz-40 Hz |  |  |  | 14020-19103 | Acoustic and Vibration | 7880 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Coniditioning amplifier | ICP mode at frequency range: 40 Hz - 5 kHz |  |  |  | 14020-19103 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
2530 | Electrical Generator | 10 mV, 100 mV, 1 V, and 10V | 0.05% for Range 10 V 0.10% for Range 1 V 0.15% for Range 100 mV 1.0% for Range 10 mV | 12,500 / piece |  | 14020-19105 | Acoustic and Vibration | 12500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Coniditioning amplifier | Voltage mode at frequency rang:40 Hz - 5 Hz |  |  |  | 14020-19120 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasound Power Meter | 1 MHz to 20 MHz 1mW to 1000 mW | 3.9 % to 13.7 % | 10,000/piece |  | 14030 - 11101 | Acoustic and Vibration | 10000 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Ultrasonic Transducer (ultrasonic transducer and generator) | 0.5 W to 10 W 1 MHz 0.5 W to 4 W 3 MHz 0.5 W to 2 W | 6.0 % 7.0% 7.0% | 12,500/piece |  | 14030 - 11102 | Acoustic and Vibration | 12500 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasonic Transducer | 1 MHz to 20 MHz 1mW to 1000 mW | 3.9 % to 13.7 % | 11,700/piece |  | 14030 - 12101 | Acoustic and Vibration | 11700 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasound Power in Medical Device (Therapeutic Ultrasound) - Onsite Service |  |  | 3,000 |  | 14030 - 15101 | Acoustic and Vibration | 3000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasound Power Meter for Physiotherap | 1 MHz |  |  |  | 14030-11101 | Acoustic and Vibration | 7880 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Ultrasound Power Calibration System | ¾ÔÊÑÂ¡ÒÃÇÑ´ÂèÒ¹¤ÇÒÁ¶Õè 1,3 áÅÐ 5 àÁ¡¡ÃÐàÎÔÃµ«ì ·Õè¢¹Ò´¡ÓÅÑ§ 0.5 ÇÑµµì |  |  |  | 14030-11101 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter
| Ultrasound Power Meter | Power:0.5 W to 10 W/ Frequency 1 MHz Power:0.5 W to 10 W/ Frequency 3 MHz Power:0.5 W to 2 W/ Frequency 5 MHz |  |  |  | 14030-11101 | Acoustic and Vibration | 7880 | Electrical Metrology Department | 7 | meter, power, phase, ac, wattmeter
| Ultrasonic Transducer | ¾ÔÊÑÂ¡ÒÃÇÑ´ÂèÒ¹¤ÇÒÁ¶Õè 1, 3 áÅÐ 5 àÁ¡¡ÐàÎÔÃµ«ì·Õè¢¹Ò´¡ÓÅÑ§ 1, 2, 4, 8 áÅÐ 10 ÇÑµµì |  |  |  | 14030-11102 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
2527 | Ultrasound Power and Time Accuracy (Therapeutic Ultrasound) - Onsite Service | 1 MHz to 3 MHz / 0.5 W to 10.0 W 1 MHz to 3 MHz / 60 s and 240 s | 8.00% 1.00% | 5,000/piece |  | 14030-11103 | Acoustic and Vibration | 5000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Needle hydrophone (free-field sensitivity - hydrophone) | Frequency : ranges: 0.5 MHz-15MHz,2 MHz,3MHz,4 MHz,5MHz,10MHz,15MHz | 13.0% to 14.0% | 24,375/piece |  | 14033 - 11101 | Acoustic and Vibration | 24375 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Immersion Ultrasonic Transducer | 1 Mhz to 15 MHz 50 kPa to <20MPa rms | 14.40% to 15.30% | 7,500/piece |  | 14033 - 13101 | Acoustic and Vibration | 7500 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasound Check Source | 100 mW to 1 W | 6% | 10,000/piece |  | 14033 - 13111 | Acoustic and Vibration | 10000 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasund transducer for physiotherapy(Planar Scanning Method) | 1 Mhz ,3 MHz 50 kPa to <20MPa rms | 17.00% | 7,500/piece |  | 14033 - 63301 | Acoustic and Vibration | 7500 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasonic transducer measurement(Effective radiation area(Aer)and the beam non-uniformity ratio(Rbn)) | Frequency : 1 MHz Frequency : 3 MHz | (Aer)7.0% (Rbn) 7.0% (f awf) 4.0% (Aer) 7.0% (Rbn)7.0% (f awf) 4.0% | 7500 THB/1 ¤ÇÒÁ¶Õè |  | 14033 - 63302 | Acoustic and Vibration | 7500 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Ultrasound Check Source | Frequency 3.5 MHz, Power : 100 mW and 1 W Uncertainty 6% |  |  |  | 14033-13111 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Effective radiation area (AER) and Beam non-uniformity ratio (RBN) | Frequency : 1 MHz |  |  |  | 14033-63302 | Acoustic and Vibration | 7880 | Acoustic and Vibration | 0 | gauge, calibration, testing, machine, laser
| Extensometer | 10%-100% | According to | 17,500 |  | 15000 - 11010 | Innovation and Measuring Devices Development Group | 17500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer | 1%-10%, 10%-100% | ISO 9513 | 22,500 |  | 15000 - 11020 | Innovation and Measuring Devices Development Group | 22500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer | 0.1%-1%, 1%-10%, 10%-100% |  | 30,000 |  | 15000 - 11030 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Gauge length for Extensometer | All sizes |  | 2,500/ea |  | 15000 - 11040 | Innovation and Measuring Devices Development Group | 2500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Jig figure | All side | According to specification | 2,500/hour |  | 15000 - 11050 | Innovation and Measuring Devices Development Group | 2500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Revers engineering | All sizes | According to Specification | 3,000/ea |  | 15000 - 11060 | Innovation and Measuring Devices Development Group | 3000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| 3D Inspection | All sizes | According to Specification | 2,500/ea |  | 15000 - 11070 | Innovation and Measuring Devices Development Group | 2500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Displacement Measuring Devices | 10% - 100% | According to CP-Inutm 1013 | 5000 |  | 15000 - 11080 | Innovation and Measuring Devices Development Group | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Extensometer Calibration Apparatus |  |  | 17,500 |  | 15000 - 11090 | Innovation and Measuring Devices Development Group | 17500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer Calibration Apparatus |  |  | 22,500 |  | 15000 - 11100 | Innovation and Measuring Devices Development Group | 22500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer Calibration Apparatus |  |  | 30,000 |  | 15000 - 11110 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CR |  |  | 18,530/set |  | 15000 - 22010 | Innovation and Measuring Devices Development Group | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CRM | 30 J ± 10 J | According to ISO148-3 | 18,530/set |  | 15000 - 22010 | Innovation and Measuring Devices Development Group | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CRM | 100 J ± 10 J | According to ISO148-3 | 18,530/set |  | 15000 - 22020 | Innovation and Measuring Devices Development Group | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CR |  |  | 18,530/set |  | 15000 - 22020 | Innovation and Measuring Devices Development Group | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for small tool (Mirometer, Vernier, Dial, Micrometer Head, Height gauge) |  |  | 30,000/ea |  | 15000 - 22030 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for small tool (Mirometer, Vernier, Dial, Micrometer Head, Height gauge) | 2 devices |  | 30,000/ea |  | 15000 - 22030 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for small tool (Mirometer, Vernier, Dial, Micrometer Head, Height gauge) |  |  | 40,000/ea |  | 15000 - 22040 | Innovation and Measuring Devices Development Group | 40000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for small tool (Mirometer, Vernier, Dial, Micrometer Head, Height gauge) | 4 devices |  | 40,000/ea |  | 15000 - 22040 | Innovation and Measuring Devices Development Group | 40000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for small tool (Mirometer, Vernier, Dial, Micrometer Head, Height gauge) |  |  | 60,000/ea |  | 15000 - 22050 | Innovation and Measuring Devices Development Group | 60000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for small tool (Mirometer, Vernier, Dial, Micrometer Head, Height gauge) | 8 devices |  | 60,000/ea |  | 15000 - 22050 | Innovation and Measuring Devices Development Group | 60000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for instrument |  |  | 30,000/ea |  | 15000 - 22060 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for instrument | 2 devices |  | 30,000/ea |  | 15000 - 22060 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for instrument |  |  | 40,000/ea |  | 15000 - 22070 | Innovation and Measuring Devices Development Group | 40000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Data link for instrument | 4 devices |  | 40,000/ea |  | 15000 - 22070 | Innovation and Measuring Devices Development Group | 40000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Software I-link for instrument |  |  | 15,000 |  | 15000 - 22080 | Innovation and Measuring Devices Development Group | 15000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Software I-link for instrument | 1 devices |  | 15,000 |  | 15000 - 22080 | Innovation and Measuring Devices Development Group | 15000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Software I-link for instrument |  |  | 35,000 |  | 15000 - 22090 | Innovation and Measuring Devices Development Group | 35000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Software I-link for instrument | 1 devices |  | 35,000 |  | 15000 - 22090 | Innovation and Measuring Devices Development Group | 35000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CRM ASTM E23 |  |  | 18,530/set |  | 15000 - 22100 | Innovation and Measuring Devices Development Group | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CRM ASTM E23 |  |  | 18,530/set |  | 15000 - 22110 | Innovation and Measuring Devices Development Group | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| V-notch Charpy Impact CRM ASTM E23 |  |  | 18,530/set |  | 15000 - 22120 | Innovation and Measuring Devices Development Group | 18530 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer | 10%-100% | According to | 17,500 |  | 15000 - 31010 | Innovation and Measuring Devices Development Group | 17500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer | 1%-10%, 10%-100% | ISO 9513 | 22,500 |  | 15000 - 31020 | Innovation and Measuring Devices Development Group | 22500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer | 0.1%-1%, 1%-10%, 10%-100% |  | 30,000 |  | 15000 - 31030 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Gauge length for Extensometer |  |  | 2,500/ea |  | 15000 - 31040 | Innovation and Measuring Devices Development Group | 2500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Callibration of metellic -charpy pendulum Testing Machine |  |  | 100,000 |  | 15000 - 31050 | Innovation and Measuring Devices Development Group | 100000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Calibration of metallic -charpy pendulum Testing Machine | Low(<40J) and high (?40J)energy | According to ISO148-2 | 40,450 |  | 15000 - 31051 | Innovation and Measuring Devices Development Group | 40450 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Calibration of metallic -charpy pendulum Testing Machine | Low(<40J) and high (?40J)energy | According to ISO148-2 | 71,500 |  | 15000 - 31052 | Innovation and Measuring Devices Development Group | 71500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Calibration of metallic -charpy pendulum Testing Machine | Low(<40J) and high (?40J)energy | According to ISO148-2 | 22,710 |  | 15000 - 31053 | Innovation and Measuring Devices Development Group | 22710 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Jig Fixture |  |  | 2,500/hour |  | 15000 - 31060 | Innovation and Measuring Devices Development Group | 2500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Reverse engineering | All sizes | According to Specification | 3,000/ea | IGES,STEP File | 15000 - 31070 | Innovation and Measuring Devices Development Group | 3000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| 3D Inspection | All sizes | According to Specification | 2,500/ea | STL File | 15000 - 31080 | Innovation and Measuring Devices Development Group | 2500 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Displacement Measuring Devices | 10%-100% | According to CP-INutm 1013 | 5,000 |  | 15000 - 31090 | Innovation and Measuring Devices Development Group | 5000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Calibration of Plastics – Verification of Charpy, Izod Impact testing Machine |  |  | 18,000 / Impact type |  | 15000 - 31100 | Innovation and Measuring Devices Development Group | 18000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Extensometer Calibration Apparatus |  |  | 17,500 |  | 15000 - 31110 | Innovation and Measuring Devices Development Group | 17500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer Calibration Apparatus |  |  | 22,500 |  | 15000 - 31120 | Innovation and Measuring Devices Development Group | 22500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Extensometer Calibration Apparatus |  |  | 30,000 |  | 15000 - 31130 | Innovation and Measuring Devices Development Group | 30000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Calibration of pendulum-type hammer impact machines for use with plastic specimens ASTM D256 (Izod |  |  | 18,000 / Impact type |  | 15000 - 31200 | Innovation and Measuring Devices Development Group | 18000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Calibration of Gauge length Calibrator |  |  | 25,000 |  | 15000 - 31300 | Innovation and Measuring Devices Development Group | 25000 | Dimensional Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Robot calibration | All sizes |  | 10,000 |  | 15000 - 31400 | Innovation and Measuring Devices Development Group | 10000 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Machine Calibration | All sizes |  | 10,000 / 1 axis |  | 15000 - 31500 | Innovation and Measuring Devices Development Group | 10000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Twin robot measurement | All sizes |  | 10,000 |  | 15000 - 31600 | Innovation and Measuring Devices Development Group | 10000 | Electrical Metrology Department | 6 | measurement, detector, additional, frequency, irradance
| Displacement Measuring Systems and Devices Used in Material Testing Machines | Displacement length up to 2,000 mm. | According toASTM E2309/E2309M-16 | 6,250 |  | 15000 - 31700 | Innovation and Measuring Devices Development Group | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Speed for material testing machines | Speed 0.05 mm/Min to 2,500 mm/Min. | According toASTM E2658-15 | 6,250 |  | 15000 - 31800 | Innovation and Measuring Devices Development Group | 6250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Callibration of metellic -charpy pendulum Testing Machine | Low(<40J) and high (?40J)energy | According to ISO148-2 | 81,250 |  | 15000 - 32010 | Innovation and Measuring Devices Development Group | 81250 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
2526 | Calibration of spring hammers | 0 - 2 J | Environmental testing-Part 2-75: Tests - Test Eh: Hammer tests According to IEC 60068-2-75 | 3,500/ea |  | 15000-12000 | Innovation and Measuring Devices Development Group | 3500 | Innovation and Measuring Devices Development Group | 0 | gauge, calibration, testing, machine, laser
| Ventilator Analyzer | Airway pressure (0 to 150) hPa Tidal volumes (30 to 600) mL Oxygen concentration (30 to 100)% | Pressure 0.15 mbar Volume 0.7% Oxygen concentration 0.55% | 7,500 |  | 16000-10000 | Medical Metrology Program | 7500 | Electrical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Pulse rate of NIBP Simulator | (30 to 200) BPM | 0.50 BPM | 3,500 |  | 16000-10001 | Medical Metrology Program | 3500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Volume of syringe | (30 to 200) BPM | 0.50 BPM | 2,500 |  | 16000-10002 | Medical Metrology Program | 2500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Infusion pump analyzer - Flow rate | 1 mL/h to 300 mL/h | 0.55%rdg | 4500/chanel+4,500 at 1 mL/h | CP-MMP 1003 | 16000-10003 | Medical Metrology Program | 4500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Infusion pump analyzer - Occlusion pressure | 100 mmHg to 1,200 mmHg | 14 mmHg | 4500/chanel+4,500 at 1 mL/h | CP-MMP 1003 | 16000-10003 | Medical Metrology Program | 4500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| Vital signs simulator - Pulse rate of NIBP | 30-300 BPM | 0.5 BPM | 7,000 | CP-MMP 1004 | 16000-10004 | Medical Metrology Program | 7000 | Medical Metrology Program | 0 | gauge, calibration, testing, machine, laser
| Vital signs simulator - NIBP simulator - staitc pressure | 0-300 mmHg | 0.5 mmHg | 7,000 | CP-MMP 1004 | 16000-10004 | Medical Metrology Program | 7000 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 30-300 BPM | 0.1 BPM | 3,500 | CP-MMP 1004 | 16000-10005 | Medical Metrology Program | 3500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 0 to 20 mV/V | 0.015% + 0.00032 mV/V | 3,500 | CP-MMP 1004 | 16000-10006 | Medical Metrology Program | 3500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 1-25 kW | 0.03% | 3,500 | CP-MMP 1004 | 16000-10007 | Medical Metrology Program | 3500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 0.2-7 mV | 0.010 mV | 3,500 | CP-MMP 1004 | 16000-10008 | Medical Metrology Program | 3500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| #NAME? | 30-300 BPM | 0.6 BPM | 3,500 | CP-MMP 1004 | 16000-10009 | Medical Metrology Program | 3500 | Mechanical Metrology Department | 0 | gauge, calibration, testing, machine, laser
| PT Sound Level Meter |  |  | 7,254 ºÒ· | ¤èÒ¨Ñ´à»ÃÕÂºà·ÕÂº 7,254 ºÒ· äÁèÃÇÁ vat | pt1 | Acoustic and Vibration | 7254 | Acoustic and Vibration | 7 | meter, power, phase, ac, wattmeter